## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | ZINNAT 250MG
Reference Substance: | Cefuroxime axetil
% age Purity: | 96.5
Weight (mg): | 28.32
Standard Conc (mg/mL): | 0.2272
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 31445676 | 6004.12 | 1.29 | 13.1
2 | 31416311 | 6008.87 | 1.3 | 13.1
3 | 31369709 | 6047.14 | 1.31 | 13.1
4 | 31408353 | 6036.2 | 1.3 | 13.11
5 | 31241363 | 6071.03 | 1.3 | 13.11
6 | 31366576 | 6224.85 | 1.32 | 13.12
Averages: | 31374664.6667 | 6065.3683 | 1.3033 | 13.1067
RSD: | 0.2288%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | ISOMER B
Reference Substance: | Cefuroxime axetil
% age Purity: | 96.5
Weight (mg): | 28.32
Standard Conc (mg/mL): | 0.2272
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 33406196 | 6348.23 | 1.32 | 15.1
2 | 33367774 | 6364.45 | 1.32 | 15.1
3 | 33314006 | 6388.56 | 1.32 | 15.1
4 | 33392004 | 6399.22 | 1.33 | 15.1
5 | 33212414 | 6443.8 | 1.33 | 15.11
6 | 33350409 | 6608.11 | 1.34 | 15.12
Averages: | 33340467.1667 | 6425.395 | 1.3267 | 15.105
RSD: | 0.2117%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | JOYFRIDA
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | ZINNAT 250MG
Laboratory Ref No: |  | NDQA201509331
Active Ingredient: |  | Cefuroxime
Label Claim: |  | Each tablet contains 250mg Cefuroxime as Cefuroxime axetil
Date Analysis Started: |  | 18th Jan 2016
Date Analysis Completed: |  | 8th Feb 2016
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 453.5 | -0.2493%
|  | 451.87 | -0.6079%
|  | 453.11 | -0.3351%
|  | 455.03 | 0.08721%
|  | 451.65 | -0.6562%
|  | 457.27 | 0.5799%
|  | 452.07 | -0.5639%
|  | 449.28 | -1.178%
|  | 450.67 | -0.8718%
|  | 453.82 | -0.1789%
|  | 452.2 | -0.5353%
|  | 455.62 | 0.217%
|  | 457.24 | 0.5733%
|  | 454.44 | -0.04256%
|  | 460.4 | 1.268%
|  | 455.32 | 0.151%
|  | 455.75 | 0.2456%
|  | 456.11 | 0.3248%
|  | 461.92 | 1.603%
|  | 455.4 | 0.1686%
| Total | 9092.67
| Average | 454.6335
| Average | % Deviation from mean
| 454.6335 | -5% | 431.9018
|  | 5% | 477.3652
| Name |  | Date | Signature
Analysed by: | JOYFRIDA
Reviewed By:

## Cefuroxime Axetil
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ZINNAT 250MG
Laboratory Ref No: | NDQA201509331 | 35
Active Ingredient: | Cefuroxime
Label Claim: | Each tablet contains 250mg Cefuroxime as Cefuroxime axetil
Date Analysis Started: | 18th Jan 2016
Date Analysis Completed: | 8th Feb 2016
Analysis Data
Reference Substance: | Cefuroxime axetil
Code: | WRS/C66-1
% age Purity: | 96.5
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 96.5
Mwt of compound in free base form: | 424.386 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 510.47 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.83136 | free base
Initial Standard dilution | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 66430122 | 70172169.6626 | 64311443 | 69130253.8524
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 65859724 | 69569640.8093 | 64268837 | 69084455.415
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 66178243 | 69906102.1711 | 64280742 | 69097252.4482
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 66156029.6667 | 69882637.5477 | 64287007.3333 | 69103987.2385
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 28.32 |  | 27.83
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 23.5442 |  | 23.1368
Standard Dilution Factor | 100 | Purity correction (mg): | 22.7202 |  | 22.327
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.2272 |  | 0.22327
|  | Desired Concetration (mg/mL): | 0.24
|  | Desired Weight as free base (mg): | 24
|  | Desired Weight as salt (mg): | 28.8682
|  | Average Normalised Response: | 69493312.3931
|  | RSD: | 0.6728%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains 250mg Cefuroxime as Cefuroxime axetil
Each Tablet contains | 250 | Cefuroxime
Average Tablet Content Weight (mg): | 454.6335
Initial Sample dilution | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 107.97 | 1 | 65077487 | 236.5905 | 94.64%
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 64323007 | 233.8476 | 93.54%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 64400517 | 234.1294 | 93.65%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 109.88 | 1 | 67730285 | 241.9546 | 96.78%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 67390868 | 240.7421 | 96.3%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 66373291 | 237.107 | 94.84%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 250 | Assay Smp C | 108.37 | 1 | 66098297 | 239.4147 | 95.77%
Desired Sample Weight (mg): | 109.112 |  |  | 2 | 65281968 | 236.4579 | 94.58%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 64739918 | 234.4945 | 93.8%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 94.88%
[6 empty] | RSD: | 1.238%
[6 empty] | n: | 9
Comment | The amount of | Cefuroxime |  | in the sample as a percentage of the stated label claim is |  | 94.88%
Analysis Data: | Determination of Active Ingredient Dissolved after |  | 15 mins
Reference Substance: | Cefuroxime axetil
Code: | WRS/C66-1
% age Purity: | 96.5
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 96.5
Mwt of compound in free base form: | 424.386 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 510.47 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.83136 | free base
Initial Standard dilution | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 0.4338 | 0.50803 | 0.49 | 0.50788
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 0.4384 | 0.51342 | 0.4905 | 0.5084
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 0.4387 | 0.51377 | 0.491 | 0.50892
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 0.43697 | 0.51174 | 0.4905 | 0.5084
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.71 |  | 22.27
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 16.3862 |  | 18.5145
Standard Dilution Factor | 1666.6667 | Purity correction (mg): | 15.8127 |  | 17.8665
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.0094876 |  | 0.01072
|  | Desired Concetration (mg/mL): | 0.011111
|  | Desired Weight as free base (mg): | 18.5185
|  | Desired Weight as salt (mg): | 22.2749
|  | Average Normalised Peak Area: | 0.51007
|  | RSD: | 0.5401%
|  | n: | 6
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 2 | 1 | 0.4127 | 202.2757 | 80.91%
1st Dilution Sample Final Volume (mL): | 50 | 2 | 0.4282 | 209.8727 | 83.95%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 0.4657 | 228.2525 | 91.3%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 0.4628 | 226.8311 | 90.73%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 0.4738 | 232.2226 | 92.89%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 0.4646 | 227.7134 | 91.09%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 88.48%
Sample Dilution Factor | 22500 |  |  | RSD: | 5.469%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment | The amount of | Cefuroxime |  | dissolved as a percentage of the stated label claim is |  | 88.48%
Analysis Data: | Determination of Active Ingredient Dissolved after |  | 45mins
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 2 | 1 | 0.5056 | 247.8086 | 99.12%
1st Dilution Sample Final Volume (mL): | 50 | 2 | 0.4979 | 244.0346 | 97.61%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 0.4873 | 238.8393 | 95.54%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 0.4979 | 244.0346 | 97.61%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 0.4734 | 232.0265 | 92.81%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 0.5071 | 248.5438 | 99.42%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 97.02%
Sample Dilution Factor | 22500 |  |  | RSD: | 2.558%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment | The amount of | Cefuroxime |  | dissolved as a percentage of the stated label claim is |  | 97.02%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0