## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Cosatrim DS Tablets
Reference Substance: | Sulphamethoxazole
% age Purity: | 99.65
Weight (mg): | 20.49
Standard Conc (mg/mL): | 0.16392
2015-10-02 12:22:45
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 128878758 | 11212.5 | 1 | 7.8
2 | 129078980 | 11206.3 | 1 | 7.8
3 | 129419325 | 11213.7 | 1 | 7.8
4 | 129453116 | 11225.3 | 1 | 7.8
5 | 129436794 | 11227 | 1 | 7.8
6 | 129462891 | 11223.9 | 1 | 7.8
Averages: | 129288310.6667 | 11218.1167 | 1 | 7.8
RSD: | 0.1921%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data
Reference Substance: | Trimethoprim
% age Purity: | 99.3
Weight (mg): | 26.55
Standard Conc (mg/mL): | 0.04248
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 12754878 | 8286.1 | 1.1 | 4.1
2 | 12768826 | 8355.2 | 1.1 | 4.1
3 | 12797469 | 8279.5 | 1.1 | 4.1
4 | 12804055 | 8362 | 1.2 | 4.1
5 | 12796383 | 8284.3 | 1.1 | 4.1
6 | 12795080 | 8298.2 | 1.1 | 4.1
Averages: | 12786115.1667 | 8310.8833 | 1.1167 | 4.1
RSD: | 0.1529%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | Cosatrim DS Tablets
Laboratory Ref No: |  | NDQA201509397
Active Ingredient: |  | Sulphamethoxazole 800 mg, Trimethoprim 160 mg
Label Claim: |  | Each tablet contains Sulfamethoxazole BP 800mg & Trimethoprim BP 160mg
Date Analysis Started: |  | 2015-10-02 12:22:45
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1059.8 | 0.4367%
|  | 1058.31 | 0.2955%
|  | 1053.73 | -0.1386%
|  | 1045.86 | -0.8844%
|  | 1053 | -0.2077%
|  | 1059.82 | 0.4386%
|  | 1055.6 | 0.03867%
|  | 1063.77 | 0.8129%
|  | 1062.35 | 0.6784%
|  | 1047.28 | -0.7498%
|  | 1043.26 | -1.131%
|  | 1042.86 | -1.169%
|  | 1058.59 | 0.322%
|  | 1051.24 | -0.3745%
|  | 1052.51 | -0.2542%
|  | 1056.94 | 0.1657%
|  | 1056.35 | 0.1097%
|  | 1057.03 | 0.1742%
|  | 1074.59 | 1.838%
|  | 1050.95 | -0.402%
| Total | 21103.84
| Average | 1055.192
| Average | % Deviation from mean
| 1055.192 | -5% | 1002.4324
|  | 5% | 1107.9516
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## SULPHAMETHOXAZOLE
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Cosatrim DS Tablets
Laboratory Ref No: | NDQA201509397 | 29
Active Ingredient: | Sulphamethoxazole 800 mg
Label Claim: | Each tablet contains Sulfamethoxazole BP 800mg & Trimethoprim BP 160mg
Date Analysis Started: | 2015-12-11 12:23
Date Analysis Completed:
Analysis Data
Reference Substance: | Sulphamethoxazole
Code: | S 12 4
% age Purity: | 99.65
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.65
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 127465401 | 124854169.6817 | 132412639 | 128695119.6014
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 127632919 | 125018255.9407 | 132190901 | 128479606.9536 |  | 2.867%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 127633743 | 125019063.0604 | 132088482 | 128380063.3938
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 127577354.3333 | 124963829.5609 | 132230674 | 128518263.3163
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 20.49 |  | 20.65
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 20.49 |  | 20.65
Standard Dilution Factor | 125 | Purity correction (mg): | 20.4183 |  | 20.5777
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.16335 |  | 0.16462
|  | Desired Concentration (mg/mL): | 0.16
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 126741046.4386
|  | RSD: | 1.539%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains Sulfamethoxazole BP 800mg & Trimethoprim BP 160mg
Each Tablet contains | 800 | Sulphamethoxazole 800 mg
Average Tablet Content Weight (mg): | 1055.192
Initial Sample dilution Volume (mL): | 200 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 2 | Assay Smp A | 1065 | 1 | 126082136 | 788.5117 | 98.56%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 125943489 | 787.6446 | 98.46%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 125887252 | 787.2929 | 98.41%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1054.65 | 1 | 124656465 | 787.2463 | 98.41%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 124593969 | 786.8516 | 98.36%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 124644917 | 787.1734 | 98.4%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1063.66 | 1 | 125299686 | 784.6055 | 98.08%
Desired Sample Weight (mg): | 1055.192 |  |  | 2 | 125207612 | 784.0289 | 98%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 125396073 | 785.2091 | 98.15%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 786.5071 | 98.31%
[5 empty] | RSD: | 0.1932% | 0.1932%
[5 empty] | n: | 9 | 9
Comment: | The content of | Sulphamethoxazole 800 mg |  | in the sample as a percentage of the stated label claim is |  | 98.31%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Sulphamethoxazole
Code: | S 12 4
% age Purity: | 99.65
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.65
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 124532956 | 139203956.8174 | 129060759 | 137839568.118
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 124669583 | 139356679.5955 | 128770372 | 137529428.7776 |  | 1.148%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 124300506 | 138944122.3061 | 128797050 | 137557921.4351
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 124501015 | 139168252.9063 | 128876060.3333 | 137642306.1103
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.95 |  | 20.88
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 19.95 |  | 20.88
Standard Dilution Factor | 125 | Purity correction: | 19.8802 |  | 20.8069
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.15904 |  | 0.16646
|  | Desired Concentration (mg/mL): | 0.17778
|  | Desired Weight as free base (mg): | 22.2222
|  | Desired Weight as salt (mg): | 22.2222
|  | Average Normalised Peak Area: | 138405279.5083
|  | RSD: | 0.6163%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 2 | 1 | 137021721 | 792.0029 | 99%
1st Dilution Sample Final Volume (mL): | 10 | 2 | 137002728 | 791.8931 | 98.99%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 133309912 | 770.5481 | 96.32%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 136212997 | 787.3283 | 98.42%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 136143794 | 786.9283 | 98.37%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 134671659 | 778.4192 | 97.3%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 784.52 | 98.06%
Sample Dilution Factor | 4500 |  | RSD: | 1.076% | 1.076%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Sulphamethoxazole 800 mg |  | dissolved as a percentage of the stated label claim is |  | 98.06%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## TRIMETHOPRIM
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Cosatrim DS Tablets
Laboratory Ref No: | NDQA201509397 | 29
Active Ingredient: | Trimethoprim 160 mg
Label Claim: | Each tablet contains Sulfamethoxazole BP 800mg & Trimethoprim BP 160mg
Date Analysis Started: | 2015-12-11 12:23
Date Analysis Completed:
Analysis Data
Reference Substance: | Trimethoprim
Code: | T14 9
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 2 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 12979716 | 9846489.2667 | 9544345 | 9635715.6739
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 12981876 | 9848127.8554 | 9532005 | 9623257.5396 |  | 1.631%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 12956221 | 9828665.8208 | 9532136 | 9623389.7937
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 12972604.3333 | 9841094.3143 | 9536162 | 9627454.3357
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 26.55 |  | 19.95
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 26.55 |  | 19.95
Standard Dilution Factor | 625 | Purity correction (mg): | 26.3641 |  | 19.8103
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.042183 |  | 0.031697
|  | Desired Concentration (mg/mL): | 0.032
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 9734274.325
|  | RSD: | 1.205%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains Sulfamethoxazole BP 800mg & Trimethoprim BP 160mg
Each Tablet contains | 160 | Trimethoprim 160 mg
Average Tablet Content Weight (mg): | 1055.192
Initial Sample dilution Volume (mL): | 200 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 2 | Assay Smp A | 1065 | 1 | 9481918 | 154.4168 | 96.51%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 9473041 | 154.2722 | 96.42%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 9476308 | 154.3254 | 96.45%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1054.65 | 1 | 9422738 | 154.9589 | 96.85%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 9423904 | 154.9781 | 96.86%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 9430960 | 155.0942 | 96.93%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1063.66 | 1 | 9465807 | 154.3486 | 96.47%
Desired Sample Weight (mg): | 1055.192 |  |  | 2 | 9465640 | 154.3459 | 96.47%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 9464932 | 154.3343 | 96.46%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 154.5638 | 96.6%
[5 empty] | RSD: | 0.2193% | 0.2193%
[5 empty] | n: | 9 | 9
Comment: | The content of | Trimethoprim 160 mg |  | in the sample as a percentage of the stated label claim is |  | 96.6%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Trimethoprim
Code: | T14 9
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 2 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 9347003 | 10917296.7441 | 9928084 | 10853900.552
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 9358646 | 10930895.7647 | 9926611 | 10852290.1914 |  | 0.6743%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 9347440 | 10917807.16 | 9927261 | 10853000.8053
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 9351029.6667 | 10921999.8896 | 9927318.6667 | 10853063.8496
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.16 |  | 20.47
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 19.16 |  | 20.47
Standard Dilution Factor | 625 | Purity correction: | 19.0259 |  | 20.3267
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.030441 |  | 0.032523
|  | Desired Concentration (mg/mL): | 0.035556
|  | Desired Weight as free base (mg): | 22.2222
|  | Desired Weight as salt (mg): | 22.2222
|  | Average Normalised Peak Area: | 10887531.8696
|  | RSD: | 0.3497%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 2 | 1 | 10688206 | 157.0708 | 98.17%
1st Dilution Sample Final Volume (mL): | 10 | 2 | 10659438 | 156.648 | 97.9%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 10777614 | 158.3847 | 98.99%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 10612145 | 155.953 | 97.47%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 10620599 | 156.0772 | 97.55%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 10463250 | 153.7649 | 96.1%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 156.3164 | 97.7%
Sample Dilution Factor | 4500 |  | RSD: | 0.9768% | 0.9768%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Trimethoprim 160 mg |  | dissolved as a percentage of the stated label claim is |  | 97.7%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## SST (DISS)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Sulfran DS Tablets
Reference Substance: | Sulphamethoxazole
% age Purity: | 99.65
Weight (mg): | 19.95
Standard Conc (mg/mL): | 0.1596
2015-10-02 10:44:06
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 124564490 | 11200.7 | 1 | 7.9
2 | 124751802 | 11185.1 | 1 | 7.9
3 | 124696532 | 11211 | 1 | 7.9
4 | 124863060 | 11212.3 | 1 | 7.9
5 | 124798906 | 11221.1 | 1 | 7.9
6 | 125066972 | 11220 | 1 | 7.9
Averages: | 124790293.6667 | 11208.3667 | 1 | 7.9
RSD: | 0.1356%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data
Reference Substance: | Trimethoprim
% age Purity: | 99.3
Weight (mg): | 19.16
Standard Conc (mg/mL): | 0.030656
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 9118617 | 8541.1 | 1.1 | 4.1
2 | 9117029 | 8629.1 | 1.2 | 4.1
3 | 9099830 | 8582.7 | 1.2 | 4.1
4 | 9108647 | 8581.9 | 1.1 | 4.1
5 | 9097392 | 8591.9 | 1.1 | 4.1
6 | 9104270 | 8599.8 | 1.1 | 4.1
Averages: | 9107630.8333 | 8587.75 | 1.1333 | 4.1
RSD: | 0.0966%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By: