## SST
HPLC System Suitability Report
Analysis Data | PYRIMETHAMINE
Sample(s) | FANSIDER TABLET
Reference Substance: | Pyrimethamine
% age Purity: | 99.7
Weight (mg): | 14.12
Standard Conc (mg/mL): | 0.16944
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 23007662 | 2711.4 | 1.6 | 9.2
2 | 23096797 | 2704.8 | 1.6 | 9.1
3 | 22982036 | 2713.6 | 1.6 | 9.1
4 | 22908790 | 2740.3 | 1.6 | 9.1
5 | 23228055 | 2715 | 1.6 | 9.1
6 | 23150928 | 2731.7 | 1.6 | 9.1
Averages: | 23062378 | 2719.4667 | 1.6 | 9.1167
RSD: | 0.5113%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | SULFADOXINE
Reference Substance: | Sulfadoxine
% age Purity: | 99.9
Weight (mg): | 19.21
Standard Conc (mg/mL): | 1.921
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 380193504 | 14242.1 | 1 | 11.8
2 | 380069507 | 14298.5 | 1 | 11.8
3 | 377735420 | 14140.6 | 1 | 11.8
4 | 377526865 | 14119.6 | 1 | 11.8
5 | 381271121 | 14282.9 | 1 | 11.8
6 | 379834923 | 14254.1 | 1 | 11.8
Averages: | 379438556.6667 | 14222.9667 | 1 | 11.8
RSD: | 0.3916%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | FANSIDER TABLET
Laboratory Ref No: |  | NDQA201511497
Active Ingredient: |  | sulfadoxine & Pyrimethamine
Label Claim: |  | sulfadoxine USP 500mg pyrimethamine USP 25mg
Date Analysis Started: |  | 2015-11-05 11:24:50
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 703.72 | -0.06192%
|  | 708.32 | 0.5913%
|  | 704.68 | 0.07442%
|  | 700.27 | -0.5519%
|  | 701.83 | -0.3303%
|  | 705.15 | 0.1412%
|  | 711.47 | 1.039%
|  | 702.41 | -0.248%
|  | 702.66 | -0.2125%
|  | 701.56 | -0.3687%
|  | 705.22 | 0.1511%
|  | 693.9 | -1.456%
|  | 708.27 | 0.5842%
|  | 708.01 | 0.5473%
|  | 704.87 | 0.1014%
-100%
-100%
-100%
-100%
-100%
| Total | 10562.34
| Average | 704.156
| Average | % Deviation from mean
| 704.156 | -5% | 668.9482
|  | 5% | 739.3638
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Pyrimethamine
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | FANSIDER TABLET
Laboratory Ref No: | NDQA201511497 | 29
Active Ingredient: | sulfadoxine & Pyrimethamine
Label Claim: | sulfadoxine USP 500mg pyrimethamine USP 25mg
Date Analysis Started: | 2016-03-08
Date Analysis Completed: | 2016-03-24
Analysis Data
Reference Substance: | Pyrimethamine
Code: | P21-2
% age Purity: | 99.7
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.7
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 23277728 | 13779374.0049 | 23376457 | 13877129.1384
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 23289179 | 13786152.4848 | 23044289 | 13679941.9328 |  | 0.4658%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 23200370 | 13733581.5283 | 22823920 | 13549122.747
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 23255759 | 13766369.3394 | 23081555.3333 | 13702064.6061
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 14.12 |  | 14.08
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 14.12 |  | 14.08
Standard Dilution Factor | 83.3333 | Purity correction (mg): | 14.0776 |  | 14.0378
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.16893 |  | 0.16845
|  | Desired Concentration (mg/mL): | 0.1
|  | Desired Weight as free base (mg): | 8.3333
|  | Desired Weight as salt (mg): | 8.3333
|  | Average Normalised Response: | 13734216.9727
|  | RSD: | 0.8132%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | sulfadoxine USP 500mg pyrimethamine USP 25mg
Each Tablet contains | 25 | sulfadoxine & Pyrimethamine
Average Tablet Content Weight (mg): | 704.156
Initial Sample dilution Volume (mL): | 250 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 701.86 | 1 | 13405917 | 24.4822 | 97.93%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 13588749 | 24.8161 | 99.26%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 13600423 | 24.8374 | 99.35%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 702.64 | 1 | 13592077 | 24.7946 | 99.18%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 13663889 | 24.9256 | 99.7%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 13680452 | 24.9559 | 99.82%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 250 | Assay Smp C | 703.83 | 1 | 13599138 | 24.7656 | 99.06%
Desired Sample Weight (mg): | 704.156 |  |  | 2 | 13639114 | 24.8384 | 99.35%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 13536084 | 24.6508 | 98.6%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 24.7852 | 99.14%
[5 empty] | RSD: | 0.5807% | 0.5807%
[5 empty] | n: | 9 | 9
Comment: | The content of | sulfadoxine & Pyrimethamine |  | in the sample as a percentage of the stated label claim is |  | 99.14%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Pyrimethamine
Code: | P21-2
% age Purity: | 99.7
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.7
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 6139871 | 3634529.0593 | 5957067 | 3536335.213
2nd Dilution Standard Transfer Volume (mL): | 5 | 2 | 6161668 | 3647431.9086 | 5964114 | 3540518.5728 |  | 3.029%
2nd Dilution Standard Final Volume (mL): | 20 | 3 |  | - | 5919928 | 3514288.1295
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 6150769.5 | 3640980.4839 | 5947036.3333 | 3530380.6384
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 14.12 |  | 14.08
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 14.12 |  | 14.08
Standard Dilution Factor | 333.3333 | Purity correction: | 14.0776 |  | 14.0378
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.042233 |  | 0.042113
|  | Desired Concentration (mg/mL): | 0.025
|  | Desired Weight as free base (mg): | 8.3333
|  | Desired Weight as salt (mg): | 8.3333
|  | Average Normalised Peak Area: | 3574620.5766
|  | RSD: | 1.722%
|  | n: | 5
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 3327711 | 23.2732 | 93.09%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 3379573 | 23.6359 | 94.54%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 3638391 | 25.446 | 101.8%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 3367843 | 23.5538 | 94.22%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 3587889 | 25.0928 | 100.4%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 3481020 | 24.3454 | 97.38%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 24.2245 | 96.9%
Sample Dilution Factor | 1000 |  | RSD: | 3.675% | 3.675%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | sulfadoxine & Pyrimethamine |  | dissolved as a percentage of the stated label claim is |  | 96.9%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Sulfadoxine
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | FANSIDER TABLET
Laboratory Ref No: | NDQA201511497 | 29
Active Ingredient: | sulfadoxine & Pyrimethamine
Label Claim: | sulfadoxine USP 500mg pyrimethamine USP 25mg
Date Analysis Started: | 2016-03-08
Date Analysis Completed: | 2016-03-24
Analysis Data
Reference Substance: | Sulfadoxine
Code: | S18-1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 10 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 377954127 | 393891160.291 | 380856483 | 391816773.6081
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 376876042 | 392767616.1325 | 378286526 | 389172858.3671 |  | 1.025%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 380033532 | 396058246.6902 | 378860258 | 389763101.2307
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 378287900.3333 | 394239007.7046 | 379334422.3333 | 390250911.0687
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 19.21 |  | 19.46
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 19.21 |  | 19.46
Standard Dilution Factor | 10 | Purity correction (mg): | 19.1908 |  | 19.4405
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 1.9191 |  | 1.9441
|  | Desired Concentration (mg/mL): | 2
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 392244959.3866
|  | RSD: | 0.658%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | sulfadoxine USP 500mg pyrimethamine USP 25mg
Each Tablet contains | 500 | sulfadoxine & Pyrimethamine
Average Tablet Content Weight (mg): | 704.156
Initial Sample dilution Volume (mL): | 250 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 701.86 | 1 | 384917884 | 492.2652 | 98.45%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 389131733 | 497.6542 | 99.53%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 389529520 | 498.1629 | 99.63%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 702.64 | 1 | 390958763 | 499.4357 | 99.89%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 392354056 | 501.2182 | 100.2%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 393185780 | 502.2807 | 100.5%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 250 | Assay Smp C | 703.83 | 1 | 390465072 | 497.9617 | 99.59%
Desired Sample Weight (mg): | 704.156 |  |  | 2 | 387667075 | 494.3934 | 98.88%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 388399668 | 495.3277 | 99.07%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 497.6333 | 99.53%
[5 empty] | RSD: | 0.6462% | 0.6462%
[5 empty] | n: | 9 | 9
Comment: | The content of | sulfadoxine & Pyrimethamine |  | in the sample as a percentage of the stated label claim is |  | 99.53%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Sulfadoxine
Code: | S18-1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 10 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 20 | 1 | 100988632 | 105246977.3261 |  | -
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 104148970 | 108540575.9742 | 102291611 | 105235359.7174 |  | 1.264%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 99958425 | 104173330.0192 | 101180961 | 104092747.4237
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 101698675.6667 | 105986961.1065 | 101736286 | 104664053.5705
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.21 |  | 19.46
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 19.21 |  | 19.46
Standard Dilution Factor | 40 | Purity correction: | 19.1908 |  | 19.4405
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.47977 |  | 0.48601
|  | Desired Concentration (mg/mL): | 0.5
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Peak Area: | 105457798.0921
|  | RSD: | 1.717%
|  | n: | 5
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 97891428 | 464.1261 | 92.83%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 99372493 | 471.1482 | 94.23%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 106978931 | 507.212 | 101.4%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 98758389 | 468.2365 | 93.65%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 105416172 | 499.8026 | 99.96%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 102860732 | 487.6867 | 97.54%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 483.0354 | 96.61%
Sample Dilution Factor | 1000 |  | RSD: | 3.709% | 3.709%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | sulfadoxine & Pyrimethamine |  | dissolved as a percentage of the stated label claim is |  | 96.61%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5