## component
MICOBIOLOGY NO. | BIOL/002/2015 | DATE RECEIVED | 2015-11-18 14:15:25
Analysis Report | Quinine Dihydrochloride Microbial Assay
Sample Name: | QUIMED INJECTION
Lab Ref No: | NDQA201511507
Active Ingredient: | Quinine Dihydrochloride
Label Claim: | Each ml contains mg of Apyrogenic
Date Test Set: | 24/11/2015
Date of Results: | 27/11/2015
Analysis Data: | Bacterial Endotoxin (LAL) Test
Lysate Sensitivity, λ (Lowest conc of std used): | 0.005 | EU/mL
Sample Endotoxin Limit (E.L): | 350 | EU/vial
Sample vol: | 2 | mL
Maximum Valid Dilution (MVD) = E.L x C | 35000
λ
Standard Information: |  | Standard Calibration Curve
Control Standard Endotoxin: | 8000 EU / vial | Regression ( R ) |  | Regression Coefficient (R2)
Reconstitution vol: | 8.0mL | 0.991 |  | 0.98208
Standard stock Concentration (EU/mL) | 1000
STANDARD DILUTIONS
Desired Standard Concetration (EU/ µL) | Initial Concentration (EU/mL) | Sample vol Pipetted (µL) | Total Vol1 (µL) | pipette2 (µL) | Total Vol2 (µL)
5 | 1000 | 100 | 2000 | 200 | 2000
0.5 | 5 | 300 | 3000
0.05 | 0.5 | 200 | 2000
0.005 | 0.05 | 200 | 2000
SAMPLE DILUTION
Sample vol Pipetted (µL) | Diluent Vol1 (µL) | pipette2 (µL) | Diluent Vol2 (µL) | pipette3 (µL) | Total Vol3 (µL)
50 | 6950 | 50 | 6950
STD Volume (µl): | 50
LAL Reagent Volume (µl): | 50
Analysis Data: | Determination of Endotoxin content in sample
Y - Ordinate (B) | 6.24
Gradient of Standard Curve (A) | -0.154
Standard curve formula: Log (Y) = A*Log (X) + B
Sample Information | Sample Volume (µL) | Sample Δt at onset OD (Y) | Log Y | Log X | Antilog X (Endotoxin Conc)
D5 | 50 | 4905 | 8.498 | -14.6624 | 4.2874e-07
D6 | 50 | 4816 | 8.4797 | -14.5435 | 4.8288e-07
|  |  | Average |  | 4.5581e-07
|  |  |  | Coefficient of Variance (RSD) : | 1.295%
|  |  | Replicates (n): |  | 2
|  |  |  | Dilution Factor : | 19600
|  |  |  | Endotoxin Concentration of Sample Solution (EU/mL) : | 0.0089339
Comment: |  | The endotoxin concentration in the sample is : | 0.017868 |  | EU/vial
Analysed by: |  | Checked by |  |  | Approved By :
ERIC, DUNCAN, SHADRACK |  | Head, BAU |  |  | Director

## C
Analysis Report | PACLITAXEL INJECTION
Sample Name: | PACLITAXEL INJECTION
Lab Ref No: | NDQD201401003
Active Ingredient: | PACLITAXEL INJECTION
Label Claim: | Endotoxin Free
Date Test Set: | 2014-03-18
Date of Results: | 2014-03-18
Analysis Data: | Bacterial Endotoxin (LAL) Test
Lysate Sensitivity, λ (Lowest conc of std used): | 0.005 | EU/mL
Sample Endotoxin Limit (E.L): | 0.67 | EU/mg
Sample reconstitution vol: |  | mL
Conc of sample solution, C (Potency): | 6
Maximum Valid Dilution (MVD) = E.L x C
λ
Standard Information: |  | Standard Calibration Curve
|  | Regression ( R ) |  | Regression Coefficient (R2)
Control Standard Endotoxin: | 7500 EU / vial | -0.999 |  | 0.998
Reconstitution vol: | 7.5mL
Standard stock Concentration (EU/mL) | 1000
STD Volume (µl): | 50
LAL Reagent Volume (µl): | 50
Analysis Data: | Determination of Endotoxin content in sample
Y - Ordinate (B) | 6.14
Gradient of Standard Curve (A) | -0.128
Standard curve formula: Log (Y) = A*Log (X) + B
Sample Information | Sample Volume (µL) | Sample Δt at onset OD (Y) | Log Y | Log X | Antilog X (Endotoxin Conc)
| 50 | 4332 | 8.3738 | -17.4514 | 2.6359e-08
| 50 | 4306 | 8.3678 | -17.4044 | 2.7629e-08
|  |  | Average |  | 2.6994e-08
|  |  |  | Coefficient of Variance (RSD) : | 0.4257%
|  |  | Replicates (n): |  | 2
|  |  |  | Dilution Factor : | 1600
|  |  |  | Endotoxin Concentration of Sample Solution (EU/mL) : | 4.319e-05
Comment: |  | The endotoxin concentration in the sample is : | 4.319e-07 |  | EU/mg
Analysed by: |  | Checked by |  |  | Approved By :
|  | Head, BAU |  |  | Director