## SST
HPLC System Suitability Report
Analysis Data | Atazanavir Assay
Sample(s)
Reference Substance: | Atazanavir sulfate
% age Purity: | 99.3
Weight (mg): | 24.43
Standard Conc (mg/mL): | 0.2443
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 62515142 | 6811 | 1.05 | 8.56
2 | 62272556 | 6816 | 1.05 | 8.57
3 | 62131987 | 6833 | 1.04 | 8.56
4 | 62168388 | 6817 | 1.05 | 8.55
5 | 62283695 | 6829 | 1.06 | 8.56
6 | 62579996 | 6824 | 1.06 | 8.59
Averages: | 62325294 | 6821.6667 | 1.0517 | 8.565
RSD: | 0.2936%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
____________________________________________________________________________________________________________________________________________________
Analysis Data | Ritonavir Assay
Reference Substance: | Ritonavir
% age Purity: | 99.3
Weight (mg): | 20.36
Standard Conc (mg/mL): | 0.2036
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 48617907 | 9610 | 1.07 | 13.21
2 | 48245930 | 9577 | 1.07 | 13.23
3 | 48257258 | 9621 | 1.06 | 13.22
4 | 48293508 | 9569 | 1.07 | 13.2
5 | 48346887 | 9589 | 1.06 | 13.21
6 | 48473087 | 9548 | 1.07 | 13.25
Averages: | 48372429.5 | 9585.6667 | 1.0667 | 13.22
RSD: | 0.3018%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
____________________________________________________________________________________________________________________________________________________
Analysis Data | Atazanavir Dissolution
Reference Substance: | Atazanavir Sulfate
% age Purity: | 99.3
Weight (mg): | 15.15
Standard Conc (mg/mL): | 0.1515
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 91670617 | 7362 | 1.1 | 5.2
2 | 91493550 | 7364 | 1.1 | 5.2
3 | 91642731 | 7341 | 1.1 | 5.2
4 | 91772523 | 7405 | 1.08 | 5.21
5 | 91723986 | 7346 | 1.1 | 5.21
6 | 91842106 | 7377 | 1.1 | 5.21
Averages: | 91690918.8333 | 7365.8333 | 1.0967 | 5.205
RSD: | 0.1311%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
____________________________________________________________________________________________________________________________________________________
Analysis Data | Atazanavir Dissolution (2nd run for S2)
Reference Substance: | Atazanavir Sulfate
% age Purity: | 99.3
Weight (mg): | 14.28
Standard Conc (mg/mL): | 0.1428
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 90314045 | 9603 | 1.06 | 9.03
2 | 90217690 | 9643 | 1.05 | 9.03
3 | 89963665 | 9614 | 1.06 | 9.03
4 | 90046512 | 9647 | 1.05 | 9.03
5 | 90166017 | 9632 | 1.06 | 9.03
6 | 90070599 | 9626 | 1.06 | 9.03
Averages: | 90129754.6667 | 9627.5 | 1.0567 | 9.03
RSD: | 0.1412%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | ATAZANAVIR SULFATE /RITONAVIR TABLETS
Laboratory Ref No: |  | NDQB201601710
Active Ingredient: |  | ATAZANAVIR SULFATE 300 mg RITONAVIR 100 mg TABLETS
Label Claim: |  | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Date Analysis Started:
Date Analysis Completed:
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1977.36 | 0.1515%
|  | 1970.43 | -0.1995%
|  | 1989.83 | 0.7831%
|  | 1969.89 | -0.2269%
|  | 1944.96 | -1.49%
|  | 1989.16 | 0.7492%
|  | 1966.25 | -0.4112%
|  | 1986.76 | 0.6276%
|  | 1961.78 | -0.6376%
|  | 1978.73 | 0.2209%
|  | 1978.13 | 0.1905%
|  | 1978.95 | 0.232%
|  | 1996.48 | 1.12%
|  | 1979.5 | 0.2599%
|  | 1959.69 | -0.7435%
|  | 1972.3 | -0.1048%
|  | 1971.47 | -0.1468%
|  | 1984.06 | 0.4908%
|  | 1969.54 | -0.2446%
|  | 1962.11 | -0.6209%
| Total | 39487.38
| Average | 1974.369
| Average | % Deviation from mean
| 1974.369 | -5% | 1875.6505
|  | 5% | 2073.0874
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## ritonavir
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ATAZANAVIR SULFATE /RITONAVIR TABLETS
Laboratory Ref No: | NDQB201601710 | 29
Active Ingredient: | RITONAVIR
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Date Analysis Started:
Date Analysis Completed:
Analysis Data
Reference Substance: | Ritonavir
Code: | R9-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 48253019 | 47733959.9198 | 50223843 | 47269054.4089
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 48255490 | 47736404.3392 | 50378828 | 47414921.2713 |  | 0.7845%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 48191070 | 47672677.307 | 50353091 | 47390698.4405
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 48233193 | 47714347.1887 | 50318587.3333 | 47358224.7069
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 20.36 |  | 21.4
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 20.36 |  | 21.4
Standard Dilution Factor | 100 | Purity correction (mg): | 20.2175 |  | 21.2502
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.20217 |  | 0.2125
|  | Desired Concentration (mg/mL): | 0.2
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 47536285.9478
|  | RSD: | 0.426%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Each Tablet contains | 100 | RITONAVIR
Average Tablet Content Weight (mg): | 1974.369
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 4 | Assay Smp A | 1946.3 | 1 | 47117746 | 100.549 | 100.5%
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 47374172 | 101.0962 | 101.1%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 46823669 | 99.9215 | 99.92%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1990.18 | 1 | 48857235 | 101.9623 | 102%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 49242795 | 102.7669 | 102.8%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 49371040 | 103.0346 | 103%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 1980.25 | 1 | 48862972 | 102.4856 | 102.5%
Desired Sample Weight (mg): | 1974.369 |  |  | 2 | 49253378 | 103.3045 | 103.3%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 48666881 | 102.0743 | 102.1%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 101.9105 | 101.9%
[5 empty] | RSD: | 1.139% | 1.139%
[5 empty] | n: | 9 | 9
Comment: | The content of | RITONAVIR |  | in the sample as a percentage of the stated label claim is |  | 101.9%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Ritonavir
Code: | R9-1
% age Purity: | 1
% Water content: | 1 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 1
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 1 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 1 | 10000 | 1 | 10000
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 1 | 10000 | 1 | 10000 |  | 0%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 1 | 10000 | 1 | 10000
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 1 | 10000 | 1 | 10000
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 1 |  | 1
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 1 |  | 1
Standard Dilution Factor | 1 | Purity correction: | 0.01 |  | 0.01
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.01 |  | 0.01
|  | Desired Concentration (mg/mL): | 100
|  | Desired Weight as free base (mg): | 100
|  | Desired Weight as salt (mg): | 100
|  | Average Normalised Peak Area: | 10000
|  | RSD: | 0%
|  | n: | 6
Medium Volume (mL): | 1 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 1 | 0.01 | 0.01%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 1 | 0.01 | 0.01%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 1 | 0.01 | 0.01%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 1 | 0.01 | 0.01%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 1 | 0.01 | 0.01%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 1 | 0.01 | 0.01%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 0.01 | 0.01%
Sample Dilution Factor | 1 |  | RSD: | 0% | 0%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | RITONAVIR |  | dissolved as a percentage of the stated label claim is |  | 0.01%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Atazanavir S2
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ATAZANAVIR SULFATE /RITONAVIR TABLETS
Laboratory Ref No: | NDQB201601710 | 29
Active Ingredient: | ATAZANAVIR
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Date Analysis Started:
Date Analysis Completed:
Analysis Data
Reference Substance: | Atazanavir sulfate
Code: | A48-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 704.9 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 802.9 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.87794 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 62216550 | 163589409.2335 | 50749945 | 166153425.2712
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 62164205 | 163451775.6356 | 50929377 | 166740878.9404
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 62114549 | 163321212.3738 | 50916527 | 166698808.5201
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 62165101.3333 | 163454132.4143 | 50865283 | 166531037.5772
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 24.43 |  | 19.62
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 21.4481 |  | 17.2252
Standard Dilution Factor | 100 | Purity correction (mg): | 21.298 |  | 17.1047
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.21298 |  | 0.17105
|  | Desired Concetration (mg/mL): | 0.56
|  | Desired Weight as free base (mg): | 56
|  | Desired Weight as salt (mg): | 63.7855
|  | Average Normalised Response: | 164992584.9957
|  | RSD: | 1.03%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Each Tablet contains | 300 | ATAZANAVIR
Average Tablet Content Weight (mg): | 1974.369
Initial Sample dilution (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 1946.3 | 1 | 172085958 | 296.2495 | 98.75%
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 172987089 | 297.8008 | 99.27%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 170902849 | 294.2127 | 98.07%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1990.18 | 1 | 180962969 | 304.6627 | 101.6%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 182210273 | 306.7626 | 102.3%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 182690907 | 307.5718 | 102.5%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 1980.25 | 1 | 176669666 | 298.9261 | 99.64%
Desired Sample Weight (mg) | 1842.7444 |  |  | 2 | 178224254 | 301.5565 | 100.5%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 176010631 | 297.811 | 99.27%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100.2%
[6 empty] | RSD: | 1.587%
[6 empty] | n: | 9
Comment | The content of | ATAZANAVIR |  | in the sample as a percentage of the stated label claim is |  | 100.2%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Atazanavir sulfate
Code: | A48-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 704.9 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 802.9 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.87794 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 93367420 | 106037305.4715 | 87516882 | 106718762.5156
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 93304567 | 105965923.3689 | 87597375 | 106816916.3021
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 93278045 | 105935802.3543 | 87353335 | 106519332.0394
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 62165101.3333 | 105979677.0649 | 50865283 | 106685003.619
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 15.15 |  | 14.11
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 13.3008 |  | 12.3878
Standard Dilution Factor | 50 | Purity correction (mg): | 13.2077 |  | 12.3011
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.26415 |  | 0.24602
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 17.0854
|  | Average Normalised Peak Area: | 106332340.342
|  | RSD: | 0.3756%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 116595646 | 328.9563 | 109.7%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 105253989 | 296.9576 | 98.99%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 110469090 | 311.6712 | 103.9%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 73046264 | 206.0886 | 68.7%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 96865741 | 273.2915 | 91.1%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 87539126 | 246.9779 | 82.33%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 92.44%
Sample Dilution Factor | 1000 |  |  | RSD: | 16.32%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Atazanavir sulfate
Code: | A48-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 704.9 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 802.9 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.87794 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 90364968 | 108879934.1672 | 95402361 | 107850069.2135
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 90267789 | 108762844.0674 | 95590372 | 108062611.1166
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 90387782 | 108907422.5498 | 95543797 | 108009959.2019
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 90340179.6667 | 108850066.9281 | 95512176.6667 | 107974213.1773
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 14.28 |  | 15.22
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 12.537 |  | 13.3623
Standard Dilution Factor | 50 | Purity correction: | 12.4493 |  | 13.2687
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.24899 |  | 0.26537
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 17.0854
|  | Average Normalised Peak Area: | 108412140.0527
|  | RSD: | 0.4494%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 61683949 | 170.6929 | 56.9%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 103095981 | 285.289 | 95.1%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 93623649 | 259.077 | 86.36%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 91787714 | 253.9966 | 84.67%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 88692158 | 245.4305 | 81.81%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 65702625 | 181.8135 | 60.6%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 77.57%
Sample Dilution Factor | 1000 |  |  | RSD: | 19.7%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 85.01%
RSD: | 19.33%
n: | 12
Comment | The amount of | ATAZANAVIR |  | dissolved as a percentage of the stated label claim is |  | 85.01%
| Name |  |  | Date |  | Signature
Analysed by: | Kefa |  |  | 26th Feb 2016
Reviewed By:
5

## RITONAVIR 2
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ATAZANAVIR SULFATE /RITONAVIR TABLETS
Laboratory Ref No: | NDQB201601710 | 24
Active Ingredient: | RITONAVIR
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Date Analysis Started:
Date Analysis Completed:
Analysis Data
Reference Substance: | Ritonavir
Code: | R9-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 48253019 | 47733959.9198 | 50223843 | 47269054.4089
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 48255490 | 47736404.3392 | 50378828 | 47414921.2713
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 48191070 | 47672677.307 | 50353091 | 47390698.4405
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 48233193 | 47714347.1887 | 50318587.3333 | 47358224.7069
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 20.36 |  | 21.4
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 20.36 |  | 21.4
Standard Dilution Factor | 100 | Purity correction (mg): | 20.2175 |  | 21.2502
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.20217 |  | 0.2125
|  | Desired Concetration (mg/mL): | 0.2
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 47536285.9478
|  | RSD: | 0.426%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Each Tablet contains | 100 | RITONAVIR
Average Tablet Content Weight (mg): | 1974.369
Initial Sample dilution | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 1946.3 | 1 | 47117746 | 100.549 | 100.5%
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 47374172 | 101.0962 | 101.1%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 46823669 | 99.9215 | 99.92%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1990.18 | 1 | 48857235 | 101.9623 | 102%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 49242795 | 102.7669 | 102.8%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 49371040 | 103.0346 | 103%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 1980.25 | 1 | 48862972 | 102.4856 | 102.5%
Desired Sample Weight (mg): | 1974.369 |  |  | 2 | 49253378 | 103.3045 | 103.3%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 48666881 | 102.0743 | 102.1%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 101.9%
[6 empty] | RSD: | 1.139%
[6 empty] | n: | 9
Comment | The amount of | RITONAVIR |  | I the sample as a percentage of the stated label claim is |  | 101.9%
Analysis Data: | Determination of Active Ingredient Dissolved after |  | 45 Minutes
Reference Substance: | Ritonavir
Code: | R9-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 53386783 | 50753445.5528 | 52180249 | 51261423.8603
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 53126407 | 50505912.7667 | 52007200 | 51091422.0242
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 53224536 | 50599201.4906 | 51933422 | 51018943.1572
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 53245908.6667 | 50619519.9367 | 52040290.3333 | 51123929.6806
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 11.77 |  | 11.39
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 11.77 |  | 11.39
Standard Dilution Factor | 100 | Purity correction (mg): | 11.6876 |  | 11.3103
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.11688 |  | 0.1131
|  | Desired Concetration (mg/mL): | 0.11111
|  | Desired Weight as free base (mg): | 11.1111
|  | Desired Weight as salt (mg): | 11.1111
|  | Average Normalised Response: | 50871724.8086
|  | RSD: | 0.5857%
|  | n: | 6
| Determination of Active Ingredient Dissolved after |  | 45 Minutes
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 26559564 | 52.2089 | 52.21%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 28806811 | 56.6264 | 56.63%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 28688034 | 56.3929 | 56.39%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 33875090 | 66.5892 | 66.59%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 32919957 | 64.7117 | 64.71%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 31533243 | 61.9858 | 61.99%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 59.75%
Sample Dilution Factor | 900 |  |  | RSD: | 9.297%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment | The amount of | RITONAVIR |  | dissolved as a percentage of the stated label claim is |  | 59.75%
Analysis Data: | Determination of Active Ingredient Dissolved after |  | 90 Minutes
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 45922178 | 90.2705 | 90.27%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 48139652 | 94.6295 | 94.63%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 47691335 | 93.7482 | 93.75%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 48875560 | 96.0761 | 96.08%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 51290747 | 100.8237 | 100.8%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 48181382 | 94.7115 | 94.71%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 95.04%
Sample Dilution Factor | 900 |  |  | RSD: | 3.621%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment | The amount of | RITONAVIR |  | dissolved as a percentage of the stated label claim is |  | 95.04%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## ATAZANAVIR 2
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ATAZANAVIR SULFATE /RITONAVIR TABLETS
Laboratory Ref No: | NDQB201601710 | 24
Active Ingredient: | ATAZANAVIR
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Date Analysis Started:
Date Analysis Completed:
Analysis Data
Reference Substance: | Atazanavir sulfate
Code: | A48-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 704.9 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 802.9 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.87794 | free base
Initial Standard dilution | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 62216550 | 163589409.2335 | 50749945 | 166153425.2712
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 62164205 | 163451775.6356 | 50929377 | 166740878.9404
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 62114549 | 163321212.3738 | 50916527 | 166698808.5201
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 62165101.3333 | 163454132.4143 | 50865283 | 166531037.5772
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 24.43 |  | 19.62
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 21.4481 |  | 17.2252
Standard Dilution Factor | 100 | Purity correction (mg): | 21.298 |  | 17.1047
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.21298 |  | 0.17105
|  | Desired Concetration (mg/mL): | 0.56
|  | Desired Weight as free base (mg): | 56
|  | Desired Weight as salt (mg): | 63.7855
|  | Average Normalised Response: | 164992584.9957
|  | RSD: | 1.03%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film coated tablet contains:Atazanavir (as sulfate) equivalent to Atazanavir 300 mg Ritonavir USP 100 mg.
Each Tablet contains | 300 | ATAZANAVIR
Average Tablet Content Weight (mg): | 1974.369
Initial Sample dilution | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 1946.3 | 1 | 172085958 | 296.2495 | 98.75%
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 172987089 | 297.8008 | 99.27%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 170902849 | 294.2127 | 98.07%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1990.18 | 1 | 180962969 | 304.6627 | 101.6%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 182210273 | 306.7626 | 102.3%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 182690907 | 307.5718 | 102.5%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 1980.25 | 1 | 176669666 | 298.9261 | 99.64%
Desired Sample Weight (mg): | 1842.7444 |  |  | 2 | 178224254 | 301.5565 | 100.5%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 176010631 | 297.811 | 99.27%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100.2%
[6 empty] | RSD: | 1.587%
[6 empty] | n: | 9
Comment | The amount of | ATAZANAVIR |  | I the sample as a percentage of the stated label claim is |  | 100.2%
Analysis Data: | Determination of Active Ingredient Dissolved after |  | 45 Minutes
Reference Substance: | Atazanavir sulfate
Code: | A48-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 704.9 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 802.9 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.87794 | free base
Initial Standard dilution | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 39579342 | 42350494.2731 | 34381284 | 42314662.2039
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 39648486 | 42424479.3984 | 34375350 | 42307358.9511
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 39544972 | 42313717.8535 | 34360361 | 42288911.2843
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 39590933.3333 | 42362897.175 | 34372331.6667 | 42303644.1464
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 16.08 |  | 13.98
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 14.1173 |  | 12.2736
Standard Dilution Factor | 50 | Purity correction (mg): | 14.0185 |  | 12.1877
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.28037 |  | 0.24375
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 17.0854
|  | Average Normalised Response: | 42333270.6607
|  | RSD: | 0.1156%
|  | n: | 6
| Determination of Active Ingredient Dissolved after |  | 45 Minutes
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 37387986 | 264.9546 | 88.32%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 36715348 | 260.1879 | 86.73%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 42459439 | 300.8941 | 100.3%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 41668084 | 295.2861 | 98.43%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 30389424 | 215.3584 | 71.79%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 33629726 | 238.3212 | 79.44%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 87.5%
Sample Dilution Factor | 1000 |  |  | RSD: | 12.48%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment | The amount of | ATAZANAVIR |  | dissolved as a percentage of the stated label claim is |  | 87.5%
Analysis Data: | Determination of Active Ingredient Dissolved after |  | 90 Minutes
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 42272373 | 299.5684 | 99.86%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 43680633 | 309.5483 | 103.2%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 42390857 | 300.4081 | 100.1%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 42961098 | 304.4492 | 101.5%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 39759464 | 281.7604 | 93.92%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 38379841 | 271.9835 | 90.66%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 98.21%
Sample Dilution Factor | 1000 |  |  | RSD: | 4.931%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment | The amount of | ATAZANAVIR |  | dissolved as a percentage of the stated label claim is |  | 98.21%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## SST 2
HPLC System Suitability Report (DISSOLUTION REPEAT)
Analysis Data | Atazanavir
Sample(s)
Reference Substance: | Atazanavir sulfate
% age Purity: | 99.3
Weight (mg): | 16.08
Standard Conc (mg/mL): | 0.3216
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 39824625 | 3290 | 0.94 | 4.12
2 | 39798905 | 3346 | 0.95 | 4.12
3 | 39787043 | 3352 | 0.94 | 4.12
4 | 39787614 | 3327 | 0.94 | 4.12
5 | 39780857 | 3360 | 0.94 | 4.12
6 | 39650703 | 3352 | 0.96 | 4.12
Averages: | 39771624.5 | 3337.8333 | 0.945 | 4.12
RSD: | 0.154%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
____________________________________________________________________________________________________________________________________________________
Analysis Data | Ritonavir
Reference Substance: | Ritonavir
% age Purity: | 99.3
Weight (mg): | 11.77
Standard Conc (mg/mL): | 0.1177
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 53214981 | 7469 | 1.04 | 7.07
2 | 53209643 | 7483 | 1.06 | 7.05
3 | 53432706 | 7493 | 1.06 | 7.07
4 | 53133704 | 7454 | 1.05 | 7.08
5 | 53035333 | 7425 | 1.05 | 7.07
6 | 53377391 | 7393 | 1.05 | 7.08
Averages: | 53233959.6667 | 7452.8333 | 1.0517 | 7.07
RSD: | 0.2793%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
____________________________________________________________________________________________________________________________________________________
| Name |  |  | Date | Signature
Analysed by:
Reviewed By: