## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Cosatrim Suspension
Reference Substance: | Sulphamethoxazole
% age Purity: | 99.58
Weight (mg): | 48.97
Standard Conc (mg/mL): | 0.19099
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 116534485 | 11443.32 | 0.86 | 7.39
2 | 116757073 | 11438.35 | 0.88 | 7.39
3 | 114572089 | 11396.59 | 0.88 | 7.39
4 | 117766233 | 11393.39 | 0.88 | 7.4
5 | 116393464 | 11384.78 | 0.88 | 7.4
6
Averages: | 116404668.8 | 11411.286 | 0.876 | 7.394
RSD: | 0.9938%
n: | 5
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
The resolution between Sulphamethoxazole and Trimethoprim should be NLT 5.0
Analysis Data | TRIMETHOPRIM
Reference Substance: | Trimethoprim
% age Purity: | 99.6
Weight (mg): | 22.49
Standard Conc (mg/mL): | 0.035862
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 7806060 | 6816.43 | 1.37 | 4.36
2 | 7816493 | 6746.02 | 1.34 | 4.36
3 | 7670809 | 6784.5 | 1.35 | 4.36
4 | 7805743 | 6782.79 | 1.34 | 4.36
5 | 7803248 | 6830.18 | 1.36 | 4.36
6
Averages: | 7780470.6 | 6791.984 | 1.352 | 4.36
RSD: | 0.7906%
n: | 5
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
The resolution between Sulphamethoxazole and Trimethoprim should be NLT 5.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Relative Density
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Relative Density Test Report
#REF!
Sample Name: | COSATRIM
Laboratory Ref No: | NDQB201602768
Active Ingredient: | Sulphamethoxazole, Trimethoprim
Label Claim: | Each 5ml contains Sulphamethoxazole 200mg and Trimethoprim 40mg
Date Analysis Started: | 15th March 2016
Date Analysis Completed: | 22nd March 2016
| Pyknometer Mass (g) | Pyknometer + Water (g) | Pyknometer + Sample (g)
| 21.9379 | 46.8203 | 50.5055
|  | 46.7742 | 50.5315
|  | 46.7854 | 50.419
| 21.9379 | 46.7933 | 50.4853
| Mass of Water (g): | 24.8554
| Mass of Sample (g): | 28.5474
| Relative Density of Sample: | 1.1485
| Name |  | Date | Signature
Analysed by:
Reviewed By:
0

## Sulphamethoxazole
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | COSATRIM
Laboratory Ref No: | NDQB201602768 | 25
Active Ingredient: | Sulphamethoxazole, Trimethoprim
Label Claim: | Each 5ml contains Sulphamethoxazole 200mg and Trimethoprim 40mg
Date Analysis Started: | 15th March 2016
Date Analysis Completed: | 22nd March 2016
Analysis Data
Reference Substance: | Sulphamethoxazole
Code: | NQCL/WRS/S12-2
% age Purity: | 99.58
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.58
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 154561691 | 129817951.3392 | 155020583 | 126822334.4564
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 155247503 | 130393971.88 | 155052578 | 126848509.565
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 154871314 | 130078006.8761 | 155263466 | 127021037.0317
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 154893502.6667 | 130096643.3651 | 155112209 | 126897293.6844
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 19.13 |  | 19.64
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 19.13 |  | 19.64
Standard Dilution Factor | 100 | Purity correction (mg): | 19.0497 |  | 19.5575
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.1905 |  | 0.19558
|  | Desired Concentration (mg/mL): | 0.16
|  | Desired Weight as free base (mg): | 16
|  | Desired Weight as salt (mg): | 16
|  | Average Normalised Response: | 128496968.5247
|  | RSD: | 1.372%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each 5ml contains Sulphamethoxazole 200mg and Trimethoprim 40mg
Each | 5 | contains | 200 | Sulphamethoxazole, Trimethoprim
Relative Density of sample: | 1.1485
Each | 5 | is equivalent to | 5.7427
Initial Sample dilution (mL): | 100 |  | Powder Weight (g) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 11.3658 | 1 | 123363396 | 194.0304 | 97.02%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 123562968 | 194.3443 | 97.17%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 123643221 | 194.4705 | 97.24%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 12.1179 | 1 | 132380566 | 195.2901 | 97.65%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 132797927 | 195.9058 | 97.95%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 133345487 | 196.7136 | 98.36%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 2500 | Assay Smp C | 10.3206 | 1 | 115948112 | 200.8363 | 100.4%
Desired Powder Weight (g): | 11.4854 |  |  | 2 | 115984543 | 200.8994 | 100.4%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 116451644 | 201.7085 | 100.9%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 98.57%
[6 empty] | RSD: | 1.588%
[6 empty] | n: | 9
Comment: | The content of | Sulphamethoxazole, Trimethoprim |  | in the sample as a percentage of the stated label claim is |  | 98.57%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## Trimethoprim
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | COSATRIM
Laboratory Ref No: | NDQB201602768 | 25
Active Ingredient: | Sulphamethoxazole, Trimethoprim
Label Claim: | Each 5ml contains Sulphamethoxazole 200mg and Trimethoprim 40mg
Date Analysis Started: | 15th March 2016
Date Analysis Completed: | 22nd March 2016
Analysis Data
Reference Substance: | Trimethoprim
Code: | NQCL/WRS/T7-2
% age Purity: | 99.6
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.6
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 10135336 | 8871632.8876 | 10110463 | 9087540.9878
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 10169598 | 8901623.002 | 10115760 | 9092302.066
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 10156616 | 8890259.6355 | 10149079 | 9122250.0296
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 10153850 | 8887838.5084 | 10125100.6667 | 9100697.6945
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 24.47 |  | 23.83
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 24.47 |  | 23.83
Standard Dilution Factor | 666.6667 | Purity correction (mg): | 24.3721 |  | 23.7347
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.036558 |  | 0.035602
|  | Desired Concentration (mg/mL): | 0.032
|  | Desired Weight as free base (mg): | 21.3333
|  | Desired Weight as salt (mg): | 21.3333
|  | Average Normalised Response: | 8994268.1014
|  | RSD: | 1.307%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each 5ml contains Sulphamethoxazole 200mg and Trimethoprim 40mg
Each | 5 | contains | 40 | Sulphamethoxazole, Trimethoprim
Relative Density of sample: | 1.1485
Each | 5 | is equivalent to | 5.7427
Initial Sample dilution (mL): | 100 |  | Powder Weight (g) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 11.3658 | 1 | 8877138 | 39.8945 | 99.74%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 8897595 | 39.9865 | 99.97%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 8874935 | 39.8846 | 99.71%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 12.1179 | 1 | 9574480 | 40.3579 | 100.9%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 9593836 | 40.4395 | 101.1%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 9612515 | 40.5182 | 101.3%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 2500 | Assay Smp C | 10.3206 | 1 | 8371144 | 41.4305 | 103.6%
Desired Powder Weight (g): | 11.4854 |  |  | 2 | 8351906 | 41.3353 | 103.3%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 8373544 | 41.4424 | 103.6%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 101.5%
[6 empty] | RSD: | 1.61%
[6 empty] | n: | 9
Comment: | The content of | Sulphamethoxazole, Trimethoprim |  | in the sample as a percentage of the stated label claim is |  | 101.5%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0