## SST LAMIVUDINE
HPLC System Suitability Report
Analysis Data | Assay
Sample(s)
Reference Substance: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
% age Purity: | NDQB201604870
Weight (mg): | Lamivudine 300mg Tablets
Standard Conc (mg/mL): | Each film-coated tablet Lamivudine USP 300mg
2016-05-06 11:16
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 19478301 | 10048.7 | 1.1 | 5.4
2 | 19532279 | 10007.9 | 1.1 | 5.4
3 | 19482165 | 10083.3 | 1.1 | 5.4
4 | 19273805 | 10051.4 | 1.1 | 5.4
5 | 19391104 | 10032.4 | 1.1 | 5.4
6 | 19520116 | 10026.3 | 1.1 | 5.4
Averages: | 19446295 | 10041.6667 | 1.1 | 5.4
RSD: | 0.5036%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 109088594 | 14142.5 | 1.1 | 5.9
2 | 107920038 | 14129.1 | 1.1 | 5.9
3 | 107586448 | 14310.2 | 1.1 | 5.9
4 | 107664730 | 14116.6 | 1.1 | 5.9
5 | 109199466 | 14170.7 | 1.1 | 5.9
6 | 109923327 | 14328 | 1.1 | 5.9
Averages: | 108563767.1667 | 14199.5167 | 1.1 | 5.9
RSD: | 0.8936%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST LAMIVUDINE (S2)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s)
Reference Substance: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
% age Purity: | NDQB201604870
Weight (mg): | Lamivudine 300mg Tablets
Standard Conc (mg/mL): | Each film-coated tablet Lamivudine USP 300mg
2016-05-06 11:16
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 19478301 | 10048.7 | 1.1 | 5.4
2 | 19532279 | 10007.9 | 1.1 | 5.4
3 | 19482165 | 10083.3 | 1.1 | 5.4
4 | 19273805 | 10051.4 | 1.1 | 5.4
5 | 19391104 | 10032.4 | 1.1 | 5.4
6 | 19520116 | 10026.3 | 1.1 | 5.4
Averages: | 19446295 | 10041.6667 | 1.1 | 5.4
RSD: | 0.5036%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 79913579 | 14025.3 | 1.1 | 5.9
2 | 79706144 | 13984.3 | 1.1 | 5.9
3 | 79752313 | 14119.6 | 1.1 | 5.9
4 | 80544326 | 14185 | 1.1 | 5.9
5 | 79872559 | 14023.3 | 1.1 | 5.9
6 | 80938458 | 14054.4 | 1.1 | 5.9
Averages: | 80121229.8333 | 14065.3167 | 1.1 | 5.9
RSD: | 0.6266%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST TENOFOVIR
HPLC System Suitability Report
Analysis Data | Assay
Sample(s)
Reference Substance: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
% age Purity: | NDQB201604870
Weight (mg): | Tenofovir Disoproxil Fumarate 300mg Tablets
Standard Conc (mg/mL): | Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
2016-05-06 11:16
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 12989431 | 168103.9 | 1.1 | 15.1
2 | 12933420 | 168953 | 1.1 | 15.1
3 | 12946994 | 169196.9 | 1.1 | 15.1
4 | 12948055 | 168815 | 1.1 | 15.1
5 | 12965307 | 168063.1 | 1.1 | 15.1
6 | 13034559 | 167938.7 | 1.1 | 15.1
Averages: | 12969627.6667 | 168511.7667 | 1.1 | 15.1
RSD: | 0.2867%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 66515987 | 153930.3 | 1.1 | 15.4
2 | 65112297 | 153318.5 | 1.2 | 15.4
3 | 64940332 | 154152.9 | 1.1 | 15.4
4 | 65119629 | 153605.4 | 1.1 | 15.4
5 | 65904855 | 153495.9 | 1.2 | 15.4
6 | 66364008 | 155618.9 | 1.1 | 15.4
Averages: | 65659518 | 154020.3167 | 1.1333 | 5.9
RSD: | 1.055%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST TENOFOVIR (S 2)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s)
Reference Substance: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
% age Purity: | NDQB201604870
Weight (mg): | Tenofovir Disoproxil Fumarate 300mg Tablets
Standard Conc (mg/mL): | Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
2016-05-06 11:16
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 12989431 | 168103.9 | 1.1 | 15.1
2 | 12933420 | 168953 | 1.1 | 15.1
3 | 12946994 | 169196.9 | 1.1 | 15.1
4 | 12948055 | 168815 | 1.1 | 15.1
5 | 12965307 | 168063.1 | 1.1 | 15.1
6 | 13034559 | 167938.7 | 1.1 | 15.1
Averages: | 12969627.6667 | 168511.7667 | 1.1 | 15.1
RSD: | 0.2867%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 60229549 | 151627.2 | 1.1 | 15.4
2 | 59735916 | 152564.6 | 1.1 | 15.4
3 | 59683412 | 151971.3 | 1.2 | 15.4
4 | 60364519 | 152966.8 | 1.1 | 15.4
5 | 59848734 | 152882.3 | 1.2 | 15.4
6 | 60643377 | 152018.4 | 1.2 | 15.4
Averages: | 60084251.1667 | 152338.4333 | 1.15 | 5.9
RSD: | 0.6445%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST EFAVIRENZ
HPLC System Suitability Report
Analysis Data | Assay
Sample(s)
Reference Substance: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
% age Purity: | NDQB201604870
Weight (mg): | Efavirenz 600mg
Standard Conc (mg/mL): | Efavirenz 600mg
2016-05-06 11:16
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 45820174 | 152797 | 1.1 | 21.9
2 | 45629637 | 153261.4 | 1 | 21.9
3 | 45688883 | 153184.5 | 1 | 21.9
4 | 45695210 | 153070.4 | 1 | 21.9
5 | 45743469 | 153260.4 | 1.1 | 21.9
6 | 45982179 | 153373.7 | 1 | 21.9
Averages: | 45759925.3333 | 153157.9 | 1.0333 | 21.9
RSD: | 0.2756%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 185235996 | 137308.6 | 1.1 | 15.4
2 | 183604383 | 137825.6 | 1.1 | 15.4
3 | 183315682 | 137962.3 | 1.1 | 15.4
4 | 183559259 | 137928.7 | 1.1 | 15.4
5 | 185884643 | 137540.5 | 1.1 | 15.4
6 | 187123195 | 136706.8 | 1.1 | 15.4
Averages: | 184787193 | 137545.4167 | 1.1 | 5.9
RSD: | 0.8361%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST EFAVIRENZ (S 2)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s)
Reference Substance: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
% age Purity: | NDQB201604870
Weight (mg): | Efavirenz 600mg
Standard Conc (mg/mL): | Efavirenz 600mg
2016-05-06 11:16
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 45820174 | 152797 | 1.1 | 21.9
2 | 45629637 | 153261.4 | 1 | 21.9
3 | 45688883 | 153184.5 | 1 | 21.9
4 | 45695210 | 153070.4 | 1 | 21.9
5 | 45743469 | 153260.4 | 1.1 | 21.9
6 | 45982179 | 153373.7 | 1 | 21.9
Averages: | 45759925.3333 | 153157.9 | 1.0333 | 21.9
RSD: | 0.2756%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 192843437 | 138077.5 | 1.1 | 21.5
2 | 193155761 | 137894.3 | 1.1 | 21.5
3 | 192979529 | 138949.8 | 1.1 | 21.5
4 | 195040305 | 138603.4 | 1.1 | 21.5
5 | 193410962 | 139101.7 | 1.1 | 21.5
6 | 195734157 | 138793.8 | 1.2 | 21.5
Averages: | 193860691.8333 | 138570.0833 | 1.1167 | 5.9
RSD: | 0.628%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | TENOFOVIR DISOPROXIL FUMARATE/ LAMIVUDINE/ EFAVIRENZ TABLETS 300 mg/300 mg /600 mg
Laboratory Ref No: |  | NDQB201604870
Active Ingredient: |  | Tenofovir Disoproxil Fumarate 300mg, Lamivudine 300mg & Efavirenz 600mg tablets
Label Claim: |  | Each tablet contains Tenofovir Disoproxil Fumarate 300mg, Lamivudine 300mg & Efavirenz 600mg tablets
Date Analysis Started: |  | 2016-04-20 11:19:08
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1925 | 1.76%
|  | 1876.54 | -0.8013%
|  | 1885.29 | -0.3387%
|  | 1915.82 | 1.275%
|  | 1892.84 | 0.06037%
|  | 1890.8 | -0.04747%
|  | 1854.57 | -1.963%
|  | 1868.81 | -1.21%
|  | 1891.84 | 0.007506%
|  | 1928.15 | 1.927%
|  | 1896.14 | 0.2348%
|  | 1922.92 | 1.65%
|  | 1850.37 | -2.185%
|  | 1908.91 | 0.9099%
|  | 1852.33 | -2.081%
|  | 1902.63 | 0.5779%
|  | 1894.49 | 0.1476%
|  | 1930.88 | 2.071%
|  | 1881.93 | -0.5164%
|  | 1863.7 | -1.48%
| Total | 37833.96
| Average | 1891.698
| Average | % Deviation from mean
| 1891.698 | -5% | 1797.1131
|  | 5% | 1986.2829
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Lamivudine
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: | NDQB201604870 | 11
Active Ingredient: | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started: | 2016-05-06 11:16
Date Analysis Completed: | 2016-05-10 11:16
Analysis Data
Reference Substance: | Lamivudine
Code: | L3-9
% age Purity: | 101.74
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 101.74
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 19226668 | 16743513.4275 | 22306491 | 16961059.3618
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 19119356 | 16650060.942 | 22233509 | 16905566.4547
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 19138415 | 16666658.442 | 22106051 | 16808651.942
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 19161479.6667 | 16686744.2705 | 22215350.3333 | 16891759.2528
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 16.93 |  | 19.39
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 16.93 |  | 19.39
Standard Dilution Factor | 250 | Purity correction (mg): | 17.2246 |  | 19.7274
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.068898 |  | 0.07891
|  | Desired Concetration (mg/mL): | 0.06
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Response: | 16789251.7617
|  | RSD: | 0.753%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Each Tablet contains | 300 | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Average Tablet Content Weight (mg): | 1891.698
Initial Sample dilution (mL): | 200 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 5 | Assay Smp A | 1889.27 | 1 | 17213938 | 307.9838 | 102.7%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 17341499 | 310.2661 | 103.4%
2nd Dilution Sample Transfer Volume (mL): | 10 |  |  | 3 | 17395668 | 311.2353 | 103.7%
2nd Dilution Sample Final Volume (mL): | 25 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1893.41 | 1 | 17282491 | 308.5343 | 102.8%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 17298614 | 308.8221 | 102.9%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 17348560 | 309.7137 | 103.2%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1890.6 | 1 | 16844459 | 301.1613 | 100.4%
Desired Sample Weight (mg) | 1891.698 |  |  | 2 | 17065428 | 305.112 | 101.7%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 16742673 | 299.3415 | 99.78%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 102.3%
[6 empty] | RSD: | 1.357%
[6 empty] | n: | 9
Comment | The content of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | in the sample as a percentage of the stated label claim is |  | 102.3%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Lamivudine
Code: | L3-9
% age Purity: | 101.74
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 101.74
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 108259694 | 82316806.1901 | 95515389 | 83179425.4862
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 108804653 | 82731173.5574 | 93477754 | 81404954.2683
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 108622756 | 82592865.5728 | 94569849 | 82356003.4722
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 108562367.6667 | 82546948.4401 | 94520997.3333 | 82313461.0756
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 19.39 |  | 16.93
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 19.39 |  | 16.93
Standard Dilution Factor | 50 | Purity correction (mg): | 19.7274 |  | 17.2246
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.39455 |  | 0.34449
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 82430204.7578
|  | RSD: | 0.7173%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 62796790 | 228.5453 | 76.18%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 86564352 | 315.046 | 105%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 75412422 | 274.4592 | 91.49%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 82513174 | 300.302 | 100.1%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 80658055 | 293.5504 | 97.85%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 77027752 | 280.3381 | 93.45%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 94.01%
Sample Dilution Factor | 1000 |  |  | RSD: | 10.61%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Lamivudine
Code: | L3-9
% age Purity: | 101.74
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 101.74
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 79038280 | 80867315.3745 | 88072417 | 82759240.6468
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 78990282 | 80818206.6464 | 87074685 | 81821699.1837
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 80280538 | 82138320.6325 | 87671577 | 82382582.2655
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 79436366.6667 | 81274614.2178 | 87606226.3333 | 82321174.032
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 14.41 |  | 15.69
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 14.41 |  | 15.69
Standard Dilution Factor | 50 | Purity correction: | 14.6607 |  | 15.963
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.29321 |  | 0.31926
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 81797894.1249
|  | RSD: | 0.9793%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 85657301 | 314.1547 | 104.7%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 86152842 | 315.9721 | 105.3%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 87820623 | 322.0888 | 107.4%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 86086757 | 315.7297 | 105.2%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 89495220 | 328.2305 | 109.4%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 87604867 | 321.2975 | 107.1%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 106.5%
Sample Dilution Factor | 1000 |  |  | RSD: | 1.664%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 100.3%
RSD: | 9.427%
n: | 12
Comment | The amount of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | dissolved as a percentage of the stated label claim is |  | 100.3%
| Name |  |  | Date |  | Signature
Analysed by: | KIPKORIR / MICHAEL
Reviewed By:
5

## Tenofovir
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: | NDQB201604870 | 11
Active Ingredient: | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started: | 2016-05-06 11:16
Date Analysis Completed: | 2016-05-10 11:16
Analysis Data
Reference Substance: | Tenofovir Disoproxil Fumarate
Code: | T11-6
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 12930095 | 12535575.3719 | 13414902 | 12358467.2625
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 12905692 | 12511916.9498 | 13416161 | 12359627.1152
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 12898653 | 12505092.7219 | 13374390 | 12321145.6163
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 12911480 | 12517528.3478 | 13401817.6667 | 12346413.3313
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 15.66 |  | 16.48
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 15.66 |  | 16.48
Standard Dilution Factor | 250 | Purity correction (mg): | 15.4721 |  | 16.2822
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.061888 |  | 0.065129
|  | Desired Concetration (mg/mL): | 0.06
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Response: | 12431970.8396
|  | RSD: | 0.7664%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Each Tablet contains | 300 | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Average Tablet Content Weight (mg): | 1891.698
Initial Sample dilution (mL): | 200 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 5 | Assay Smp A | 1889.27 | 1 | 11751035 | 283.9325 | 94.64%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 11805932 | 285.259 | 95.09%
2nd Dilution Sample Transfer Volume (mL): | 10 |  |  | 3 | 11719165 | 283.1625 | 94.39%
2nd Dilution Sample Final Volume (mL): | 25 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1893.41 | 1 | 11895054 | 286.7839 | 95.59%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 11872287 | 286.235 | 95.41%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 11864233 | 286.0409 | 95.35%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1890.6 | 1 | 11890155 | 287.0919 | 95.7%
Desired Sample Weight (mg) | 1891.698 |  |  | 2 | 11885052 | 286.9687 | 95.66%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 11876178 | 286.7544 | 95.58%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 95.27%
[6 empty] | RSD: | 0.4932%
[6 empty] | n: | 9
Comment | The content of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | in the sample as a percentage of the stated label claim is |  | 95.27%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Tenofovir Disoproxil Fumarate
Code: | T11-6
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 65776679 | 60596710.5877 | 63491314 | 61554083.8724
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 65882628 | 60694316.0155 | 62322385 | 60420820.9239
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 65759914 | 60581265.8455 | 63169453 | 61242043.4098
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 65806407 | 60624097.4829 | 62994384 | 61072316.0687
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 16.48 |  | 15.66
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 16.48 |  | 15.66
Standard Dilution Factor | 50 | Purity correction (mg): | 16.2822 |  | 15.4721
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.32564 |  | 0.30944
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 60848206.7758
|  | RSD: | 0.7328%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 44599515 | 219.8891 | 73.3%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 58825399 | 290.0269 | 96.68%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 54408378 | 268.2497 | 89.42%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 58491003 | 288.3783 | 96.13%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 57949180 | 285.7069 | 95.24%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 55109167 | 271.7048 | 90.57%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 90.22%
Sample Dilution Factor | 1000 |  |  | RSD: | 9.775%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Tenofovir Disoproxil Fumarate
Code: | T11-6
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 59453951 | 60702149.057 | 67234151 | 61306991.1004
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 59329174 | 60574752.4429 | 66687044 | 60808115.4029
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 60200461 | 61464331.5618 | 67215688 | 61290155.7428
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 59661195.3333 | 60913744.3539 | 67045627.6667 | 61135087.4154
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 14.87 |  | 16.65
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 14.87 |  | 16.65
Standard Dilution Factor | 50 | Purity correction: | 14.6916 |  | 16.4502
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.29383 |  | 0.329
|  | Desired Concetration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 61024415.8846
|  | RSD: | 0.6118%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 63002429 | 309.724 | 103.2%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 59305784 | 291.5511 | 97.18%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 60909476 | 299.4349 | 99.81%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 59396191 | 291.9955 | 97.33%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 61728126 | 303.4595 | 101.2%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 61141325 | 300.5747 | 100.2%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 99.82%
Sample Dilution Factor | 1000 |  |  | RSD: | 2.318%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 95.02%
RSD: | 8.348%
n: | 12
Comment | The amount of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | dissolved as a percentage of the stated label claim is |  | 95.02%
| Name |  |  | Date |  | Signature
Analysed by: | KIPKORIR/MICHAEL
Reviewed By:
5

## Efavirenz
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: | NDQB201604870 | 11
Active Ingredient: | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started: | 2016-05-06 11:16
Date Analysis Completed: | 2016-05-10 11:16
Analysis Data
Reference Substance: | Efavirenz
Code: | E15-3
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 45662253 | 47002530.1471 | 42965470 | 47013779.5426
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 45597828 | 46936214.1464 | 42942842 | 46989019.4783
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 45584053 | 46922034.823 | 42867707 | 46906805.0785
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 45614711.3333 | 46953593.0388 | 42925339.6667 | 46969868.0331
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 29.35 |  | 27.61
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 29.35 |  | 27.61
Standard Dilution Factor | 250 | Purity correction (mg): | 29.1445 |  | 27.4167
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.11658 |  | 0.10967
|  | Desired Concetration (mg/mL): | 0.12
|  | Desired Weight as free base (mg): | 30
|  | Desired Weight as salt (mg): | 30
|  | Average Normalised Response: | 46961730.536
|  | RSD: | 0.09694%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Each Tablet contains | 600 | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Average Tablet Content Weight (mg): | 1891.698
Initial Sample dilution (mL): | 200 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 5 | Assay Smp A | 1889.27 | 1 | 43871455 | 561.2379 | 93.54%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 44040443 | 563.3997 | 93.9%
2nd Dilution Sample Transfer Volume (mL): | 10 |  |  | 3 | 43702748 | 559.0796 | 93.18%
2nd Dilution Sample Final Volume (mL): | 25 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1893.41 | 1 | 44683090 | 570.3711 | 95.06%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 44578612 | 569.0374 | 94.84%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 44535415 | 568.486 | 94.75%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1890.6 | 1 | 45130962 | 576.9443 | 96.16%
Desired Sample Weight (mg) | 1891.698 |  |  | 2 | 45073424 | 576.2088 | 96.03%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 45077869 | 576.2656 | 96.04%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 94.83%
[6 empty] | RSD: | 1.179%
[6 empty] | n: | 9
Comment | The content of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | in the sample as a percentage of the stated label claim is |  | 94.83%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Efavirenz
Code: | E15-3
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 185856098 | 203367904.9252 | 197025667 | 202808758.7559
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 186186620 | 203729569.5001 | 193873415 | 199563981.9452
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 186050291 | 203580395.255 | 196455270 | 202221619.4795
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 186031003 | 203559289.8934 | 195784784 | 201531453.3935
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 27.61 |  | 29.35
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 27.61 |  | 29.35
Standard Dilution Factor | 50 | Purity correction (mg): | 27.4167 |  | 29.1445
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.54833 |  | 0.58289
|  | Desired Concetration (mg/mL): | 0.6
|  | Desired Weight as free base (mg): | 30
|  | Desired Weight as salt (mg): | 30
|  | Average Normalised Peak Area: | 202545371.6435
|  | RSD: | 0.7716%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 151310323 | 448.2265 | 74.7%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 194102790 | 574.9905 | 95.83%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 183674940 | 544.1001 | 90.68%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 194483787 | 576.1192 | 96.02%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 194069276 | 574.8913 | 95.82%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 180042816 | 533.3407 | 88.89%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 90.32%
Sample Dilution Factor | 1000 |  |  | RSD: | 9.117%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Efavirenz
Code: | E15-3
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 192948217 | 197802893.413 | 208521888 | 198230173.8348
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 192843768 | 197695816.4225 | 206650608 | 196451251.9037
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 195413429 | 200330131.4153 | 208687019 | 198387154.7022
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 193735138 | 198609613.7503 | 207953171.6667 | 197689526.8136
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 29.47 |  | 31.78
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 29.47 |  | 31.78
Standard Dilution Factor | 50 | Purity correction: | 29.2637 |  | 31.5575
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.58527 |  | 0.63115
|  | Desired Concetration (mg/mL): | 0.6
|  | Desired Weight as free base (mg): | 30
|  | Desired Weight as salt (mg): | 30
|  | Average Normalised Peak Area: | 198149570.2819
|  | RSD: | 0.6395%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 192511175 | 582.9269 | 97.15%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 202179206 | 612.2018 | 102%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 201413660 | 609.8837 | 101.6%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 202336150 | 612.677 | 102.1%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 199481448 | 604.0329 | 100.7%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 200525614 | 607.1947 | 101.2%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 100.8%
Sample Dilution Factor | 1000 |  |  | RSD: | 1.852%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 95.56%
RSD: | 8.263%
n: | 12
Comment | The amount of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | dissolved as a percentage of the stated label claim is |  | 95.56%
| Name |  |  | Date |  | Signature
Analysed by: | KIPKORIR/ MICHAEL
Reviewed By:
5