## SST Sulfamethoxazole
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | CO-TRI TABLETS
Reference Substance: | Sulfamethoxazole
% age Purity: | 99.28
Weight (mg): | 16.23
Standard Conc (mg/mL): | 0.1623
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Resolution(USP) | Retention Times (min)
1 | 49875935 | 10951.1 | 1 | 13.4728 | 5.9
2 | 49918100 | 10964.5 | 1 | 13.4788 | 5.9
3 | 49957335 | 10937 | 1 | 13.451 | 5.9
4 | 49851115 | 10953.6 | 1 | 13.4494 | 5.9
5 | 50015695 | 10906.8 | 1 | 13.4384 | 5.9
6 | 49997118 | 11002.2 | 1 | 13.4627 | 5.9
Averages: | 49935883 | 10952.5333 | 1 | 13.46 | 5.9
RSD: | 0.1319%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Resolution between Sulfamethoxazole and Trimethoprim is NLT 5.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) |  | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Resolution between Sulfamethoxazole and Trimethoprim is NLT 5.0
| Name |  |  |  | Date | Signature
Analysed by: | RUTTO KENNEDY |  |  |  | 19/12/2016
Reviewed By:

## SST Trimethoprim
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | CO-TRI TABLETS
Reference Substance: | Trimethoprim
% age Purity: | 99.3
Weight (mg): | 19.18
Standard Conc (mg/mL): | 0.030688
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3591472 | 7957.4 | 1.1 | 3.4
2 | 3591443 | 7960.5 | 1.1 | 3.4
3 | 3592607 | 7973.9 | 1 | 3.4
4 | 3583759 | 7955.9 | 1 | 3.4
5 | 3592894 | 7932 | 1 | 3.4
6 | 3591056 | 7956.8 | 1.1 | 3.4
Averages: | 3590538.5 | 7956.0833 | 1.05 | 3.4
RSD: | 0.09465%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO KENNEDY |  |  | 19/12/2016
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | CO-TRIMOXAZOLE TABLETS
Laboratory Ref No: |  | NDQB201612286
Active Ingredient: |  | Sulfamethoxazole & Trimethoprim
Label Claim: |  | Each tablet contains Sulfamethoxazole 400 mg and Trimethoprim 80 mg
Date Analysis Started: |  | 2016-12-15 08:08:51
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1044.72 | 0.3933%
|  | 1050.14 | 0.9142%
|  | 1038.82 | -0.1736%
|  | 1056.51 | 1.526%
|  | 1028.6 | -1.156%
|  | 1050.1 | 0.9103%
|  | 1039.98 | -0.06217%
|  | 1055.06 | 1.387%
|  | 1030.54 | -0.9693%
|  | 1024.75 | -1.526%
|  | 1045.23 | 0.4423%
|  | 1025.94 | -1.411%
|  | 1048.07 | 0.7152%
|  | 1038.18 | -0.2351%
|  | 1027.36 | -1.275%
|  | 1037.37 | -0.313%
|  | 1039.92 | -0.06794%
|  | 1054.2 | 1.304%
|  | 1040.52 | -0.01028%
|  | 1036.53 | -0.3937%
| Total | 20812.54
| Average | 1040.627
| Average | % Deviation from mean
| 1040.627 | -5% | 988.5956
|  | 5% | 1092.6583
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Sulfamethoxazole
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | CO-TRIMOXAZOLE TABLETS
Laboratory Ref No: | NDQB201612286 | 1
Active Ingredient: | Sulfamethoxazole
Label Claim: | Each tablet contains Sulfamethoxazole 400 mg and Trimethoprim 80 mg
Date Analysis Started: | 2016-12-14
Date Analysis Completed: | 2016-12-16
Analysis Data
Reference Substance: | Sulfamethoxazole
Code: | NDQE201607046
% age Purity: | 99.28
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.28
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 50184865 | 49832474.6555 | 48151933 | 50522022.1008
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 49955358 | 49604579.2181 | 48294927 | 50672054.4168 |  | 1.684%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 49901476 | 49551075.5691 | 48078312 | 50444777.3973
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 50013899.6667 | 49662709.8142 | 48175057.3333 | 50546284.6383
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 16.23 |  | 15.36
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 16.23 |  | 15.36
Standard Dilution Factor | 100 | Purity correction (mg): | 16.1131 |  | 15.2494
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.16113 |  | 0.15249
|  | Desired Concentration (mg/mL): | 0.16
|  | Desired Weight as free base (mg): | 16
|  | Desired Weight as salt (mg): | 16
|  | Average Normalised Response: | 50104497.2262
|  | RSD: | 0.9949%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains Sulfamethoxazole 400 mg and Trimethoprim 80 mg
Each Tablet contains | 800 | Sulfamethoxazole
Average Tablet Content Weight (mg): | 1040.627
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 217.09 | 1 | 51477067 | 787.9748 | 98.4969
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 51313516 | 785.4713 | 98.1839
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 51486491 | 788.1191 | 98.5149
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 201 | 1 | 48093055 | 795.1053 | 99.3882
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 47941920 | 792.6067 | 99.0758
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 48157676 | 796.1737 | 99.5217
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 211.17 | 1 | 50124311 | 788.7776 | 98.5972
Desired Sample Weight (mg): | 208.1254 |  |  | 2 | 50271701 | 791.097 | 98.8871
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 50168177 | 789.4679 | 98.6835
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 790.5326 | 98.8166
[5 empty] | RSD: | 0.4462% | 0.4462%
[5 empty] | n: | 9 | 9
Comment: | The content of | Sulfamethoxazole |  | in the sample as a percentage of the stated label claim is |  | 9882%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Sulfamethoxazole
Code: | NDQE201607046
% age Purity: | 99.28
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.28
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 50184865 | 55369416.2839 | 48151933 | 56135580.112
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 49955358 | 55116199.1312 | 48294927 | 56302282.6853 |  | 1.684%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 49901476 | 55056750.6323 | 48078312 | 56049752.6636
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 50013899.6667 | 55180788.6825 | 48175057.3333 | 56162538.487
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 16.23 |  | 15.36
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 16.23 |  | 15.36
Standard Dilution Factor | 100 | Purity correction: | 16.1131 |  | 15.2494
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.16113 |  | 0.15249
|  | Desired Concentration (mg/mL): | 0.17778
|  | Desired Weight as free base (mg): | 17.7778
|  | Desired Weight as salt (mg): | 17.7778
|  | Average Normalised Peak Area: | 55671663.5847
|  | RSD: | 0.9949%
|  | n: | 6
Medium Volume (mL): | 900 | Unit No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 51669465 | 742.4885 | 92.8111
1st Dilution Sample Final Volume (mL): | 50 | 2 | 51788969 | 744.2058 | 93.0257
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 51556493 | 740.8651 | 92.6081
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 51577194 | 741.1626 | 92.6453
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 51645151 | 742.1391 | 92.7674
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 51659889 | 742.3509 | 92.7939
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 742.202 | 92.7753
Sample Dilution Factor | 4500 |  | RSD: | 0.1593% | 0.1593%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Range
|  |  | Minimum: | 740.8651 | 92.6081
|  |  | Maximum: | 744.2058 | 93.0257
Comment: | The amount of | Sulfamethoxazole |  | dissolved as a percentage of the stated label claim is |  | 92.7753
| Range: | Minimum | 92.6081 | Maximum | 93.0257
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## Trimethoprim
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | CO-TRIMOXAZOLE TABLETS
Laboratory Ref No: | NDQB201612286 | 1
Active Ingredient: | Trimethoprim
Label Claim: | Each tablet contains Sulfamethoxazole 400 mg and Trimethoprim 80 mg
Date Analysis Started: | 2016-12-14
Date Analysis Completed: | 2016-12-16
Analysis Data
Reference Substance: | Trimethoprim
Code: | T7-4
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 3605034 | 3785659.1553 | 3803092 | 3810847.0738
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 3587484 | 3767229.8372 | 3816217 | 3823998.8376 |  | 1.18%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 3586029 | 3765701.9365 | 3803428 | 3811183.7589
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 3592849 | 3772863.643 | 3807579 | 3815343.2235
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 19.18 |  | 20.1
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 19.18 |  | 20.1
Standard Dilution Factor | 625 | Purity correction (mg): | 19.0457 |  | 19.9593
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.030473 |  | 0.031935
|  | Desired Concentration (mg/mL): | 0.032
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 3794103.4332
|  | RSD: | 0.6527%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains Sulfamethoxazole 400 mg and Trimethoprim 80 mg
Each Tablet contains | 160 | Trimethoprim
Average Tablet Content Weight (mg): | 1040.627
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 217.09 | 1 | 3900686 | 157.702 | 98.5637
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 3891170 | 157.3172 | 98.3233
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3904369 | 157.8509 | 98.6568
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 201 | 1 | 3650542 | 159.4033 | 99.627
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3639446 | 158.9187 | 99.3242
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3652463 | 159.4871 | 99.6795
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 211.17 | 1 | 3842164 | 159.6907 | 99.8067
Desired Sample Weight (mg): | 208.1254 |  |  | 2 | 3853952 | 160.1806 | 100.1129
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 3847045 | 159.8936 | 99.9335
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 158.9382 | 99.3364
[5 empty] | RSD: | 0.6628% | 0.6628%
[5 empty] | n: | 9 | 9
Comment: | The content of | Trimethoprim |  | in the sample as a percentage of the stated label claim is |  | 9934%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Trimethoprim
Code: | T7-4
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 3605034 | 4206287.9503 | 3803092 | 4234274.5264
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 3587484 | 4185810.9302 | 3816217 | 4248887.5974 |  | 1.18%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 3586029 | 4184113.2628 | 3803428 | 4234648.6211
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 3592849 | 4192070.7144 | 3807579 | 4239270.2483
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.18 |  | 20.1
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 19.18 |  | 20.1
Standard Dilution Factor | 625 | Purity correction: | 19.0457 |  | 19.9593
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.030473 |  | 0.031935
|  | Desired Concentration (mg/mL): | 0.035556
|  | Desired Weight as free base (mg): | 22.2222
|  | Desired Weight as salt (mg): | 22.2222
|  | Average Normalised Peak Area: | 4215670.4814
|  | RSD: | 0.6527%
|  | n: | 6
Medium Volume (mL): | 900 | Unit No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 3904922 | 148.206 | 92.6287
1st Dilution Sample Final Volume (mL): | 50 | 2 | 3912571 | 148.4963 | 92.8102
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 3895502 | 147.8484 | 92.4053
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 3896230 | 147.8761 | 92.4225
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 3901419 | 148.073 | 92.5456
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 3901468 | 148.0749 | 92.5468
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 148.0958 | 92.5599
Sample Dilution Factor | 4500 |  | RSD: | 0.1606% | 0.1606%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Range
|  |  | Minimum: | 147.8484 | 92.4053
|  |  | Maximum: | 148.4963 | 92.8102
Comment: | The amount of | Trimethoprim |  | dissolved as a percentage of the stated label claim is |  | 92.5599
| Range: | Minimum | 92.4053 | Maximum | 92.8102
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0