## SST Sulf
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | SULFRAN -DS TABLETS
Reference Substance: | Sulfamethoxazole
% age Purity: | 99.8
Weight (mg): | 16.47
Standard Conc (mg/mL): | 0.1647
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 65963255 | 8920.7 | 1.3 | 7.6
2 | 65960782 | 9089.8 | 1.2 | 7.6
3 | 66177968 | 8971.7 | 1.2 | 7.6
4 | 66039574 | 8476.9 | 1.2 | 7.6
5 | 66167258 | 8407.9 | 1.3 | 7.6
6 | 66432927 | 8255.4 | 1.3 | 7.6
Averages: | 66123627.3333 | 8687.0667 | 1.25 | 7
RSD: | 0.2704%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Sulfamethoxazole
% age Purity: | 99.02
Weight (mg): | 15.33
Standard Conc (mg/mL): | 0.1533
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 56970761 | 11481.7 | 1 | 6.9
2 | 57025806 | 11640.2 | 0.9 | 6.9
3 | 57263011 | 11518.8 | 1 | 6.9
4 | 57086193 | 11569.7 | 1 | 6.9
5 | 57194038 | 11714.2 | 0.9 | 6.9
6 | 57576534 | 11600.9 | 1 | 6.9
Averages: | 57186057.1667 | 11587.5833 | 0.96667 | 6.9
RSD: | 0.3835%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST trim
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | SULFRAN -DS TABLETS
Reference Substance: | trimethoprim
% age Purity: | 99.3
Weight (mg): | 19.46
Standard Conc (mg/mL): | 0.031136
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4593822 | 5507.7 | 1.6 | 3.9
2 | 4586513 | 5800.3 | 1.4 | 3.9
3 | 4592903 | 5558.8 | 1.4 | 3.9
4 | 4583015 | 5765.6 | 1.4 | 3.9
5 | 4586942 | 5620.9 | 1.4 | 3.9
6 | 4608309 | 5559.5 | 1.4 | 3.9
Averages: | 4591917.3333 | 5635.4667 | 1.4333 | 3.9
RSD: | 0.1964%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | TRIMETHOPRIM
% age Purity: | 99.3
Weight (mg): | 19.96
Standard Conc (mg/mL): | 0.031936
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4455491 | 8516.9 | 1.1 | 4.1
2 | 4444967 | 8539.6 | 1.1 | 4.1
3 | 4452746 | 8533.6 | 1.1 | 4.1
4 | 4423938 | 8521.2 | 1.1 | 4.1
5 | 4419696 | 8606.6 | 1.1 | 4.1
6 | 4439373 | 8595 | 1.1 | 4.1
Averages: | 4439368.5 | 8552.15 | 1.1 | 4.1
RSD: | 0.3334%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | SULFRAN-DS TABLETS
Laboratory Ref No: |  | NDQB201703355
Active Ingredient: |  | Sulfamethoxazole BP & Trimethoprim BP
Label Claim: |  | Each tablet contains: Sulphamethoxazole BP 800 mg and Trimethoprim BP 160 mg.
Date Analysis Started: |  | 2017-03-22 07:22:34
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1035.59 | -0.801%
|  | 1035.41 | -0.8183%
|  | 1048.76 | 0.4605%
|  | 1044.77 | 0.07831%
|  | 1049.78 | 0.5582%
|  | 1046.13 | 0.2086%
|  | 1052.12 | 0.7824%
|  | 1048.12 | 0.3992%
|  | 1034.46 | -0.9093%
|  | 1045.24 | 0.1233%
|  | 1042.46 | -0.143%
|  | 1048.65 | 0.45%
|  | 1039.54 | -0.4227%
|  | 1040.39 | -0.3413%
|  | 1038.44 | -0.528%
|  | 1049.33 | 0.5151%
|  | 1049.82 | 0.562%
|  | 1045.83 | 0.1798%
|  | 1044.07 | 0.01126%
|  | 1040.14 | -0.3652%
| Total | 20879.05
| Average | 1043.9525
| Average | % Deviation from mean
| 1043.9525 | -5% | 991.7549
|  | 5% | 1096.1501
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## SULFAMETHOXAZOLE (2)
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | SULFRAN -DS TABLETS
Laboratory Ref No: | NDQB201703355 | 1
Active Ingredient: | Sulfamethoxazole BP
Label Claim: | Each tablet contains: Sulphamethoxazole B.P. 800 mg and Trimethoprim B.P. 160 mg.
Date Analysis Started: | 2017-03-22 07:07:57
Date Analysis Completed:
Analysis Data
Reference Substance: | Sulfamethoxazole
Code: | S 12 4
% age Purity: | 99.8
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 67227040 | 65439478.8362 | 66542715 | 64852101.7719
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 67232180 | 65444482.1641 | 66640106 | 64947018.4137 |  | 0.7285%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 67476723 | 65682522.7869 | 67036722 | 65333557.8147
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 67311981 | 65522161.2624 | 66739847.6667 | 65044226.0001
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 16.47 |  | 16.45
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 16.47 |  | 16.45
Standard Dilution Factor | 100 | Purity correction (mg): | 16.4371 |  | 16.4171
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.16437 |  | 0.16417
|  | Desired Concentration (mg/mL): | 0.16
|  | Desired Weight as free base (mg): | 16
|  | Desired Weight as salt (mg): | 16
|  | Average Normalised Response: | 65283193.6313
|  | RSD: | 0.4898%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains: Sulphamethoxazole B.P. 800 mg and Trimethoprim B.P. 160 mg.
Each Tablet contains | 800 | Sulfamethoxazole BP
Average Tablet Content Weight (mg): | 1043.9525
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 2 | Assay Smp A | 204.94 | 1 | 63529029 | 793.1308 | 99.1413
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 63989383 | 798.8781 | 99.8598
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 64983699 | 811.2917 | 101.4115
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 208.64 | 1 | 63956803 | 784.3114 | 98.0389
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 64265318 | 788.0947 | 98.5118
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 64369353 | 789.3705 | 98.6713
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 211.96 | 1 | 65146606 | 786.3886 | 98.2986
Desired Sample Weight (mg): | 208.7905 |  |  | 2 | 65693945 | 792.9956 | 99.1244
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 65615423 | 792.0477 | 99.006
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 792.9455 | 99.1182
[5 empty] | RSD: | 1.023% | 1.023%
[5 empty] | n: | 9 | 9
Comment: | The content of | Sulfamethoxazole |  | in the sample as a percentage of the stated label claim is |  | 9912%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Sulfamethoxazole
Code: | S 12 6
% age Purity: | 99.02
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.02
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 57545122 | 67393950.0195 | 64189980 | 67002860.5919
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 57520834 | 67365505.1365 | 64011049 | 66816088.6246 |  | 0.6763%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 57400537 | 67224619.346 | 64137057 | 66947618.4437
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 57488831 | 67328024.834 | 64112695.3333 | 66922189.2201
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 15.33 |  | 17.2
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 15.33 |  | 17.2
Standard Dilution Factor | 100 | Purity correction: | 15.1798 |  | 17.0314
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.1518 |  | 0.17031
|  | Desired Concentration (mg/mL): | 0.17778
|  | Desired Weight as free base (mg): | 17.7778
|  | Desired Weight as salt (mg): | 17.7778
|  | Average Normalised Peak Area: | 67125107.027
|  | RSD: | 0.3537%
|  | n: | 6
Medium Volume (mL): | 900 | Unit No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 4 | 1 | 66516210 | 792.7431 | 99.0929
1st Dilution Sample Final Volume (mL): | 20 | 2 | 69589377 | 829.3693 | 103.6712
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 69524771 | 828.5993 | 103.5749
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 66667065 | 794.541 | 99.3176
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 66494041 | 792.4789 | 99.0599
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 69928656 | 833.4128 | 104.1766
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 811.8574 | 101.4822
Sample Dilution Factor | 4500 |  | RSD: | 2.52% | 2.52%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Range
|  |  | Minimum: | 792.4789 | 99.0599
|  |  | Maximum: | 833.4128 | 104.1766
Comment: | The amount of | Sulfamethoxazole |  | dissolved as a percentage of the stated label claim is |  | 101.4822
| Range: | Minimum | 99.0599 | Maximum | 104.1766
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## TRIMETHOPRIM (2)
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | SULFRAN -DS TABLETS
Laboratory Ref No: | NDQB201703355 | 1
Active Ingredient: | Trimethoprim BP
Label Claim: | Each tablet contains: Sulphamethoxazole B.P. 800 mg and Trimethoprim B.P. 160 mg.
Date Analysis Started: | 2017-03-22 07:07:57
Date Analysis Completed:
Analysis Data
Reference Substance: | TRIMETHOPRIM
Code: | T7 4
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4647576 | 4810214.1506 | 5003708 | 4780816.761
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4636907 | 4799171.7977 | 5004921 | 4781975.7276 |  | 0.4454%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4658421 | 4821438.6626 | 5032892 | 4808700.7535
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4647634.6667 | 4810274.8703 | 5013840.3333 | 4790497.7474
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 19.46 |  | 21.08
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 19.46 |  | 21.08
Standard Dilution Factor | 625 | Purity correction (mg): | 19.3238 |  | 20.9324
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.030918 |  | 0.033492
|  | Desired Concentration (mg/mL): | 0.032
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 4800386.3088
|  | RSD: | 0.34%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains: Sulphamethoxazole B.P. 800 mg and Trimethoprim B.P. 160 mg.
Each Tablet contains | 160 | Trimethoprim BP
Average Tablet Content Weight (mg): | 1043.9525
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 2 | Assay Smp A | 204.94 | 1 | 4599009 | 156.168 | 97.605
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 4633268 | 157.3313 | 98.3321
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 4689286 | 159.2335 | 99.521
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 208.64 | 1 | 4640997 | 154.799 | 96.7494
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 4665346 | 155.6112 | 97.257
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 4665393 | 155.6128 | 97.258
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 211.96 | 1 | 4669576 | 153.3127 | 95.8204
Desired Sample Weight (mg): | 208.7905 |  |  | 2 | 4716924 | 154.8672 | 96.792
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 4697820 | 154.24 | 96.4
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 155.6862 | 97.3039
[5 empty] | RSD: | 1.129% | 1.129%
[5 empty] | n: | 9 | 9
Comment: | The content of | TRIMETHOPRIM |  | in the sample as a percentage of the stated label claim is |  | 9730%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | TRIMETHOPRIM
Code: | T7 4
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4425240 | 4961517.5299 | 4821982 | 4896122.0173
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4411647 | 4946277.2473 | 4795768 | 4869504.9659 |  | 1.462%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4388944 | 4920822.9596 | 4792916 | 4866609.1152
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4408610.3333 | 4942872.5789 | 4803555.3333 | 4877412.0328
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.96 |  | 22.04
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 19.96 |  | 22.04
Standard Dilution Factor | 625 | Purity correction: | 19.8203 |  | 21.8857
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.031712 |  | 0.035017
|  | Desired Concentration (mg/mL): | 0.035556
|  | Desired Weight as free base (mg): | 22.2222
|  | Desired Weight as salt (mg): | 22.2222
|  | Average Normalised Peak Area: | 4910142.3059
|  | RSD: | 0.8045%
|  | n: | 6
Medium Volume (mL): | 900 | Unit No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 4 | 1 | 4790675 | 156.1071 | 97.5669
1st Dilution Sample Final Volume (mL): | 20 | 2 | 5045821 | 164.4212 | 102.7632
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 5040552 | 164.2495 | 102.6559
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4806508 | 156.623 | 97.8894
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 4800587 | 156.4301 | 97.7688
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 5083251 | 165.6409 | 103.5255
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 160.5786 | 100.3616
Sample Dilution Factor | 4500 |  | RSD: | 2.877% | 2.877%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Range
|  |  | Minimum: | 156.1071 | 97.5669
|  |  | Maximum: | 165.6409 | 103.5255
Comment: | The amount of | TRIMETHOPRIM |  | dissolved as a percentage of the stated label claim is |  | 100.3616
| Range: | Minimum | 97.5669 | Maximum | 103.5255
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0