## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: |  | NDQD201504198
Active Ingredient: |  | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: |  | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started: |  | 2015-04-22 14:13:59
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1756.83 | -0.418%
|  | 1743.2 | -1.191%
|  | 1763.82 | -0.02179%
|  | 1766.58 | 0.1346%
|  | 1769.84 | 0.3194%
|  | 1757.38 | -0.3868%
|  | 1764.3 | 0.005413%
|  | 1744.25 | -1.131%
|  | 1762.29 | -0.1085%
|  | 1765.75 | 0.0876%
|  | 1774.72 | 0.596%
|  | 1794.05 | 1.692%
|  | 1762.24 | -0.1114%
|  | 1778.48 | 0.8092%
|  | 1769.13 | 0.2792%
|  | 1753.3 | -0.6181%
|  | 1759.22 | -0.2825%
|  | 1758.99 | -0.2956%
|  | 1779.17 | 0.8483%
|  | 1760.55 | -0.2071%
| Total | 35284.09
| Average | 1764.2045
| Average | % Deviation from mean
| 1764.2045 | -5% | 1675.9943
|  | 5% | 1852.4147
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## SST(EFV)
HPLC System Suitability Report
Analysis Data | Assay/Dissolution
Sample(s) | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Reference Substance: | EFAVIRENZ
% age Purity: | 99.3
Weight (mg): | 29.76
Standard Conc (mg/mL): | 0.12
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 38776826 | 233868.8 | 1.28 | 19.83
2 | 38959088 | 238933.2 | 1.34 | 19.83
3 | 38779357 | 236745.9 | 1.33 | 19.83
4 | 38639337 | 229866.1 | 1.31 | 19.83
5 | 38632610 | 231240.7 | 1.29 | 19.83
6 | 38685433 | 231240.7 | 1.29 | 19.83
Averages: | 38745441.8333 | 233649.2333 | 1.3067 | 19.83
RSD: | 0.3169%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST(TEN)
HPLC System Suitability Report
Analysis Data | Assay/Dissolution
Sample(s) | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Reference Substance: | TENOFOVIR DISOPROXIL FUMARATE
% age Purity: | 99.3
Weight (mg): | 17.31
Standard Conc (mg/mL): | 0.06
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 13635499 | 158991.3 | 1.21 | 15.29
2 | 13635499 | 158991.3 | 1.21 | 15.29
3 | 13689534 | 166402.3 | 1.21 | 15.3
4 | 13671911 | 162918.1 | 1.23 | 15.3
5 | 13696323 | 165012.3 | 1.26 | 15.3
6 | 13688054 | 16311.1 | 1.23 | 15.31
Averages: | 13669470 | 138104.4 | 1.225 | 15.2983
RSD: | 0.2012%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST(LAM)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Reference Substance: | LAMIVUDIN
% age Purity: | 99.3
Weight (mg): | 17.72
Standard Conc (mg/mL): | 0.06
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 20471649 | 32151.8 | 1.3 | 6.97
2 | 20485943 | 33452.8 | 1.29 | 6.98
3 | 20417345 | 33653.4 | 1.31 | 6.97
4 | 20360029 | 34166.9 | 1.35 | 6.97
5 | 20447662 | 331708 | 1.33 | 6.97
6 | 20378160 | 331991 | 1.29 | 6.98
Averages: | 20426798 | 132853.9833 | 1.3117 | 6.9733
RSD: | 0.2483%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Efavirenz
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: | NDQD201504198 | 1
Active Ingredient: | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started: | 2015-04-22 14:13:59
Date Analysis Completed:
Analysis Data
Reference Substance: | EFAVIRENZ
Code: | ELT301VSP19320813(M)
% age Purity: | 99.3
% Water content: |  | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 38700455 | 39287568.4225 | 35150171 | 39639369.2014
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 38574570 | 39159773.6575 | 35175331 | 39667742.5066 |  | 1.164%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 38440150 | 39023314.4106 | 35187864 | 39681876.1566
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 38571725 | 39156885.4969 | 35171122 | 39662995.9549
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 29.76 |  | 26.79
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 29.76 |  | 26.79
Standard Dilution Factor | 250 | Purity correction (mg): | 29.5517 |  | 26.6025
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.11821 |  | 0.10641
|  | Desired Concentration (mg/mL): | 0.12
|  | Desired Weight as free base (mg): | 30
|  | Desired Weight as salt (mg): | 30
|  | Average Normalised Response: | 39409940.7259
|  | RSD: | 0.7355%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Each Tablet contains | 600 | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Average Tablet Content Weight (mg): | 1764.2045
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 688.74 | 1 | 72525636 | 565.6667 | 94.28%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 72546155 | 565.8268 | 94.3%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 72281873 | 563.7655 | 93.96%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 719.43 | 1 | 76847373 | 573.8057 | 95.63%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 76457479 | 570.8944 | 95.15%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 76274114 | 569.5253 | 94.92%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 761.7 | 1 | 72419447 | 510.735
Desired Sample Weight (mg): | 352.8409 |  |  | 2 | 80596156 | 568.4009 | 94.73%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 80179476 | 565.4623 | 94.24%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 94.65%
[6 empty] | RSD: | 0.5917%
[6 empty] | n: | 8
Comment: | The content of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | in the sample as a percentage of the stated label claim is |  | 94.65%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Efavirenz
Code: | ELT301VSP19320813CM
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 20 | 1 | 170219813 | 172802168.6077 | 157811017 | 177965824.602
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 170273567 | 172856738.0941 | 158597014 | 178852205.0772 |  | 3.196%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 169583735 | 172156440.8521 | 159068804 | 179384249.658
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 170025705 | 172605115.8513 | 158492278.3333 | 178734093.1124
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 29.76 |  | 26.79
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 29.76 |  | 26.79
Standard Dilution Factor | 50 | Purity correction: | 29.5517 |  | 26.6025
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.59103 |  | 0.53205
|  | Desired Concentration (mg/mL): | 0.6
|  | Desired Weight as free base (mg): | 30
|  | Desired Weight as salt (mg): | 30
|  | Average Normalised Peak Area: | 175669604.4818
|  | RSD: | 1.933%
|  | n: | 6
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 158037078 | 539.7761 | 89.96%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 159067566 | 543.2957 | 90.55%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 165467253 | 565.1538 | 94.19%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 155616690 | 531.5092 | 88.58%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 152801978 | 521.8956 | 86.98%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 158677364 | 541.963 | 90.33%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 90.1%
Sample Dilution Factor | 1000 |  |  | RSD: | 2.674%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | dissolved as a percentage of the stated label claim is |  | 90.1%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Lamivudine
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: | NDQD201504198 | 1
Active Ingredient: | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started:
Date Analysis Completed: | 21/8/2015
Analysis Data
Reference Substance: | Lamuvidine
Code: | Prs/L3-3
% age Purity: | 99.3
% Water content: |  | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 20334587 | 17334592.9975 | 22122729 | 17216908.6231
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 20239605 | 17253623.8432 | 22080059 | 17183700.8986 |  | 0.3944%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 20183041 | 17205404.8202 | 22109364 | 17206507.3754
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 20252411 | 17264540.5536 | 22104050.6667 | 17202372.299
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 17.72 |  | 19.41
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 17.72 |  | 19.41
Standard Dilution Factor | 250 | Purity correction (mg): | 17.596 |  | 19.2741
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.070384 |  | 0.077097
|  | Desired Concentration (mg/mL): | 0.06
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Response: | 17233456.4263
|  | RSD: | 0.3167%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Each Tablet contains | 300 | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Average Tablet Content Weight (mg): | 1764.2045
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 688.74 | 1 | 34096664 | 304.0776 | 101.4%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 34215013 | 305.133 | 101.7%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 35928476 | 320.4138
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 719.43 | 1 | 34748415 | 296.6704 | 98.89%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 34658066 | 295.899 | 98.63%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 34544758 | 294.9317 | 98.31%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 761.7 | 1 | 36827454 | 296.972 | 98.99%
Desired Sample Weight (mg): | 352.8409 |  |  | 2 | 36764328 | 296.463 | 98.82%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 36552194 | 294.7524 | 98.25%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 99.37%
[6 empty] | RSD: | 1.373%
[6 empty] | n: | 8
Comment: | The content of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | in the sample as a percentage of the stated label claim is |  | 99.37%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Lamuvidine
Code: | Prs/L3-3
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 20 | 1 | 100910731 | 86023210.1573 | 111421882 | 86713549.7166
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 100950491 | 86057104.3012 | 110039174 | 85637463.7921 |  | 0.1803%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 100565897 | 85729250.0665 | 110354440 | 85882818.0572
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 100809039.6667 | 85936521.5083 | 110605165.3333 | 86077943.8553
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 17.72 |  | 19.41
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 17.72 |  | 19.41
Standard Dilution Factor | 50 | Purity correction: | 17.596 |  | 19.2741
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.35192 |  | 0.38548
|  | Desired Concentration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 86007232.6818
|  | RSD: | 0.4446%
|  | n: | 6
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 88024302 | 307.0357 | 102.3%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 86707582 | 302.4429 | 100.8%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 80810474 | 281.8733 | 93.96%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 88481273 | 308.6296 | 102.9%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 85256730 | 297.3822 | 99.13%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 86355242 | 301.2139 | 100.4%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 99.92%
Sample Dilution Factor | 1000 |  |  | RSD: | 3.222%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | dissolved as a percentage of the stated label claim is |  | 99.92%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Tenofovir Didoproxil Fumerate
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EFAVIRENZ 600MG, LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: | NDQD201504198 | 1
Active Ingredient: | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Date Analysis Started:
Date Analysis Completed:
Analysis Data
Reference Substance: | Tenofovir Disoproxil Fumerate
Code: | Wrs T11-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 13668144 | 11927639.054 | 15177804 | 11763569.202
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 13598558 | 11866914.153 | 15073525 | 11682747.6791 |  | 1.507%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 13459463 | 11745531.5458 | 14990051 | 11618051.0882
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 13575388.3333 | 11846694.9176 | 15080460 | 11688122.6564
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 17.31 |  | 19.49
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 17.31 |  | 19.49
Standard Dilution Factor | 250 | Purity correction (mg): | 17.1888 |  | 19.3536
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.068755 |  | 0.077414
|  | Desired Concentration (mg/mL): | 0.06
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Response: | 11767408.787
|  | RSD: | 0.973%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film-coated tablet contains Efavirenz 600mg, Lamivudine USP 300mg, Tenofovir Disoproxil Fumarate 300mg euivalent to tenofovir disoproxil 245mg
Each Tablet contains | 300 | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets
Average Tablet Content Weight (mg): | 1764.2045
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 688.74 | 1 | 23243586 | 303.5758 | 101.2%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 23338090 | 304.8101 | 101.6%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 23486330 | 306.7462 | 102.2%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 719.43 | 1 | 23886107 | 298.6594 | 99.55%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 23860256 | 298.3362 | 99.45%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 23863221 | 298.3732 | 99.46%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 761.7 | 1 | 25339083 | 299.2446 | 99.75%
Desired Sample Weight (mg): | 352.8409 |  |  | 2 | 25404036 | 300.0116 | 100%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 25317616 | 298.991 | 99.66%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100.3%
[6 empty] | RSD: | 1.062%
[6 empty] | n: | 9
Comment: | The content of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | in the sample as a percentage of the stated label claim is |  | 100.3%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Tenofovir Disoproxil Fumerate
Code: | Wrs T11-1
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 20 | 1 | 68417627 | 59705308.9128 | 76562476 | 59339808.6245
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 68322492 | 59622288.428 | 75895897 | 58823176.0342 |  | 1.032%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 68189723 | 59506426.2664 | 76165082 | 59031808.085
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 68309947.3333 | 59611341.2024 | 76207818.3333 | 59064930.9146
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 17.31 |  | 19.49
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 17.31 |  | 19.49
Standard Dilution Factor | 50 | Purity correction: | 17.1888 |  | 19.3536
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.34378 |  | 0.38707
|  | Desired Concentration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 59338136.0585
|  | RSD: | 0.5852%
|  | n: | 6
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 59581233 | 301.229 | 100.4%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 58966905 | 298.1231 | 99.37%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 55280276 | 279.4844 | 93.16%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 60635964 | 306.5615 | 102.2%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 57704471 | 291.7406 | 97.25%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 58176843 | 294.1288 | 98.04%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 98.4%
Sample Dilution Factor | 1000 |  |  | RSD: | 3.156%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Efavirenz 600mg, Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg Tablets |  | dissolved as a percentage of the stated label claim is |  | 98.4%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Sheet1
|  | 12 | 20
|  | 350 | 10+2
475 | 500
9.5833 | 850
0.3
0.4 | 0.2