## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | SUMAFIX 50MG
Reference Substance: | Sumatriptan succinate
% age Purity: | 99.7
Weight (mg): | 9.84
Standard Conc (mg/mL): | 0.0984
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 11677017 | 5252.27 | 1 | 2.84
2 | 11648684 | 5236.71 | 1.01 | 2.84
3 | 11670624 | 5274.82 | 1.03 | 2.84
4 | 11668544 | 5239.12 | 1.02 | 2.84
5 | 11672924 | 5328.97 | 1.04 | 2.84
6 | 11660947 | 5283.96 | 1.02 | 2.84
Averages: | 11666456.6667 | 5269.3083 | 1.02 | 2.84
RSD: | 0.08752%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is NMT 1.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is NMT 1.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO/LORNA |  |  | 17/08/2016
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | SUMAFIX 50MG
Laboratory Ref No: |  | NDQD201504211
Active Ingredient: |  | Sumatriptan succinatew equivalent to sumatriptan 50mg
Label Claim: |  | each tablets contaians sumatriptan succinate equivalent to sumatriptan 50mg
Date Analysis Started: |  | 2015-04-24 11:44:10
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 170.12 | -0.4599%
|  | 171.93 | 0.5992%
|  | 172.06 | 0.6752%
|  | 170.9 | -0.003511%
|  | 171.29 | 0.2247%
|  | 172.97 | 1.208%
|  | 168.45 | -1.437%
|  | 172.84 | 1.132%
|  | 168.82 | -1.221%
|  | 168.31 | -1.519%
|  | 170.26 | -0.378%
|  | 168.64 | -1.326%
|  | 170.16 | -0.4365%
|  | 171.75 | 0.4938%
|  | 171.18 | 0.1603%
|  | 172.57 | 0.9736%
|  | 170.74 | -0.09713%
|  | 171.09 | 0.1077%
|  | 171.94 | 0.605%
|  | 172.1 | 0.6986%
| Total | 3418.12
| Average | 170.906
| Average | % Deviation from mean
| 170.906 | -7.5% | 158.0881
|  | 7.5% | 183.724
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Sumatriptan
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | SUMAFIX 50MG
Laboratory Ref No: | NDQD201504211 | 29
Active Ingredient: | Sumatriptan
Label Claim: | Each uncoated tablet contains:Sumatriptan Succinate Ph.Eur. equivalent to Sumatriptan 50mg
Date Analysis Started: | 2016-08-12
Date Analysis Completed: | 2016-08-15
Analysis Data
Reference Substance: | Sumatriptan succinate
Code: | S39-1
% age Purity: | 99.7
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.7
Mwt of compound in free base form: | 295.4 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 413.5 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.71439 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 11643798 | 16613817.2411 | 13417748 | 16971745.7434
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 11658808 | 16635234.0843 | 13409003 | 16960684.43 |  | 2.382%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 11648990 | 16621225.3856 | 13433027 | 16991071.7364
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 11650532 | 16623425.5703 | 13419926 | 16974500.6366
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 9.84 |  | 11.1
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 7.0296 |  | 7.9297
Standard Dilution Factor | 100 | Purity correction (mg): | 7.0085 |  | 7.9059
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.070085 |  | 0.079059
|  | Desired Concentration (mg/mL): | 0.1
|  | Desired Weight as free base (mg): | 10
|  | Desired Weight as salt (mg): | 13.998
|  | Average Normalised Response: | 16798963.1035
|  | RSD: | 1.147%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each uncoated tablet contains:Sumatriptan Succinate Ph.Eur. equivalent to Sumatriptan 50mg
Each Tablet contains | 50 | Sumatriptan
Average Tablet Content Weight (mg): | 170.906
Initial Sample dilution Volume (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 10 | Assay Smp A | 174.51 | 1 | 16204611 | 47.2349 | 94.47%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 16096478 | 46.9197 | 93.84%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 16194984 | 47.2069 | 94.41%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 166.32 | 1 | 15346777 | 46.9372 | 93.87%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 15245911 | 46.6287 | 93.26%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 15238350 | 46.6056 | 93.21%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 170.97 | 1 | 15891844 | 47.2824 | 94.56%
Desired Sample Weight (mg): | 170.906 |  |  | 2 | 15876119 | 47.2356 | 94.47%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 15783945 | 46.9613 | 93.92%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 47.0014 | 94%
[5 empty] | RSD: | 0.552% | 0.552%
[5 empty] | n: | 9 | 9
Comment: | The content of | Sumatriptan |  | in the sample as a percentage of the stated label claim is |  | 94%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Sumatriptan succinate
Code: | S39-1
% age Purity: | 99.7
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.7
Mwt of compound in free base form: | 295.4 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 413.5 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.71439 | free base
Initial Standard dilution volume (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 0.27 | 0.47178 | 0.305 | 0.48472
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 0.274 | 0.47877 | 0.31 | 0.49267 |  | 2.508%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 0.275 | 0.48051 | 0.306 | 0.48631
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 0.273 | 0.47702 | 0.307 | 0.4879
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 11.16 |  | 12.27
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 7.9726 |  | 8.7656
Standard Dilution Factor | 500 | Purity correction: | 7.9487 |  | 8.7393
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.015897 |  | 0.017479
|  | Desired Concentration (mg/mL): | 0.027778
|  | Desired Weight as free base (mg): | 13.8889
|  | Desired Weight as salt (mg): | 19.4416
|  | Average Normalised Peak Area: | 0.48246
|  | RSD: | 1.482%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 0.452 | 46.8433 | 93.69%
1st Dilution Sample Final Volume (mL): | 20 | 2 | 0.449 | 46.5324 | 93.06%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 0.456 | 47.2578 | 94.52%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 0.451 | 46.7396 | 93.48%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 0.452 | 46.8433 | 93.69%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 0.457 | 47.3615 | 94.72%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 46.9296 | 93.86%
Sample Dilution Factor | 1800 |  | RSD: | 0.6759% | 0.6759%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Sumatriptan |  | dissolved as a percentage of the stated label claim is |  | 93.86%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Weight Var.
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | SUMAFIX 50MG
Laboratory Ref No: | NDQD201504211 | 18
Active Ingredient: | Sumatriptan
Label Claim: | Each uncoated tablet contains:Sumatriptan Succinate Ph. Eur. Equivalent to Sumatriptan 50 mg
Date Analysis Started: | 2016-08-12
Date Analysis Completed: | 2016-08-15
Analysis Data | Determination of weight variation of the Sample
Label Claim: | Each uncoated tablet contains:Sumatriptan Succinate Ph. Eur. Equivalent to Sumatriptan 50 mg
Each Tablet/Capsule/vial contains | 94 | Sumatriptan | Please enter the percentage amount determined from the Assay test
| Tablet/Capsule/Vial No. | Weight: | Content as % of the average content
| 1 | 171.8 | 94.1869 |  | Calculation of acceptance value (AV)
| 2 | 171.2 | 93.858 |  | Average | 94
| 3 | 172.01 | 94.3021 |  | k | 2.4
| 4 | 172.64 | 94.6475 |  | s | 0.76578
| 5 | 172.9 | 94.79 |  | M | 98.5
| 6 | 171.96 | 94.2747 |  | AV (≤ 15) | 6.3379
| 7 | 170.18 | 93.2988
| 8 | 168.69 | 92.4819
| 9 | 170.19 | 93.3043
| 10 | 173.02 | 94.8558
| Average tablet: | 171.459 | 94
| RSD: | 0.8147% | 0.8147%
| n: | 10 | 10
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
Each Tablet contains
Average Tablet Content Weight (mg): | 170.906
0