## SST (AH)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Arteunate & Amodiaquine Tablets
Reference Substance: | Amodiaquine HCl
% age Purity: | 99.9
Weight (mg): | 12.2
Standard Conc (mg/mL): | 0.1952
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 29203065 | 2850.7 | 1.9 | 2
2 | 29476354 | 2840.5 | 1.9 | 2
3 | 29358341 | 2905.9 | 1.8 | 2
4 | 29400085 | 2855.3 | 1.9 | 2
5 | 29225054 | 2872.1 | 1.9 | 2
6 | 29533270 | 2855.9 | 1.9 | 2
Averages: | 29366028.1667 | 2863.4 | 1.8833 | 2
RSD: | 0.4512%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Amodiaquine HCl
% age Purity: | 99.9
Weight (mg): | 12.2
Standard Conc (mg/mL): | 0.05856
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 29203065 | 2850.7 | 1.9 | 2
2 | 29476354 | 2840.5 | 1.9 | 2
3 | 29358341 | 2905.9 | 1.8 | 2
4 | 29400085 | 2855.3 | 1.9 | 2
5 | 29225054 | 2872.1 | 1.9 | 2
6 | 29533270 | 2855.9 | 1.9 | 2
Averages: | 29366028.1667 | 2863.4 | 1.8833 | 2
RSD: | 0.4512%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:

## SST (A)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Arteunate & Amodiaquine Tablets
Reference Substance: | Artesunate
% age Purity: | 99.3
Weight (mg): | 22.9
Standard Conc (mg/mL): | 4.58
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3363557 | 8390.3 | 1 | 8.2
2 | 3364730 | 8341.1 | 1 | 8.2
3 | 3368981 | 8353.8 | 1 | 8.2
4 | 3372122 | 8373.4 | 1 | 8.2
5 | 3373427 | 8397.2 | 1 | 8.2
6 | 3376361 | 8390.5 | 1 | 8.2
Averages: | 3369863 | 8374.3833 | 1 | 8.2
RSD: | 0.1495%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Artesunate
% age Purity: | 99.3
Weight (mg): | 22.9
Standard Conc (mg/mL): | 0.0458
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3343174 | 8183 | 1 | 8.2
2 | 3347139 | 8177.6 | 1 | 8.2
3 | 3354260 | 8219 | 1 | 8.2
4 | 3352541 | 8244.6 | 1 | 8.2
5 | 3353202 | 8234.4 | 1 | 8.2
6 | 3351001 | 8229.4 | 1 | 8.2
Averages: | 3350219.5 | 8214.6667 | 1 | 8.2
RSD: | 0.1269%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | Arteunate & Amodiaquine Tablets
Laboratory Ref No: |  | NDQD201508213
Active Ingredient: |  | Arteunate & Amodiaquine
Label Claim: |  | Each bilayered tablet contains Artesunate 25 mg & Amodiaquine HCl 67.5mg
Date Analysis Started: |  | 2015-11-04
Date Analysis Completed: |  | 2015-11-30
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 174.51 | -0.5658%
|  | 176.18 | 0.3857%
|  | 172.52 | -1.7%
|  | 173.08 | -1.381%
|  | 172.46 | -1.734%
|  | 174.05 | -0.8279%
|  | 169.98 | -3.147%
|  | 175.83 | 0.1863%
|  | 175.73 | 0.1293%
|  | 179.16 | 2.084%
|  | 173.1 | -1.369%
|  | 175.66 | 0.08946%
|  | 178.35 | 1.622%
|  | 176.78 | 0.7276%
|  | 174.83 | -0.3835%
|  | 177.92 | 1.377%
|  | 173.43 | -1.181%
|  | 180.57 | 2.887%
|  | 179.62 | 2.346%
|  | 176.3 | 0.4541%
| Total | 3510.06
| Average | 175.503
| Average | % Deviation from mean
| 175.503 | -7.5% | 162.3403
|  | 7.5% | 188.6657
| Name |  | Date | Signature
Analysed by: | Bugigi |  | 2015-12-01
Reviewed By:

## Amodiaquine HCl
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Artesunate & Amodiaquine Tablets
Laboratory Ref No: | NDQD201508213 | 21
Active Ingredient: | Amodiaquine
Label Claim: | Each bilayered tablet contains Artesunate 25 mg & Amodiaquine HCl 67.5mg
Date Analysis Started: | 2015-11-04
Date Analysis Completed: | 2015-11-30
Analysis Data
Reference Substance: | Amodiaquine HCl
Code: | A7 1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 355.79 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 428.79 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.82975 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 29814137 | 36851727.2255 | 27428250 | 36123352.4411
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 30041651 | 37132945.6242 | 27902280 | 36747655.9514 |  | 0.8459%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 30142039 | 37257030.0877 | 28373383 | 37368104.5657
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 29999275.6667 | 37080567.6458 | 27901304.3333 | 36746370.9861
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 12.2 |  | 11.45
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 10.123 |  | 9.5007
Standard Dilution Factor | 62.5 | Purity correction (mg): | 10.1129 |  | 9.4912
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.16181 |  | 0.15186
|  | Desired Concentration (mg/mL): | 0.2
|  | Desired Weight as free base (mg): | 12.5
|  | Desired Weight as salt (mg): | 15.0647
|  | Average Normalised Response: | 36913469.3159
|  | RSD: | 1.229%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each bilayered tablet contains Artesunate 25 mg & Amodiaquine HCl 67.5mg
Each Tablet contains | 67.5 | Amodiaquine
Average Tablet Content Weight (mg): | 175.503
Initial Sample dilution Volume (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 4 | Assay Smp A | 174.04 | 1 | 41771184 | 71.3194 | 105.7%
1st Dilution Sample Final Volume (mL): | 25 |  |  | 2 | 41414578 | 70.7105 | 104.8%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 42125309 | 71.924 | 106.6%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 172.35 | 1 | 38720648 | 66.7592 | 98.9%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 38592980 | 66.5391 | 98.58%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 39088502 | 67.3934 | 99.84%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 312.5 | Assay Smp C | 174.42 | 1 | 39674033 | 67.5911 | 100.1%
Desired Sample Weight (mg): | 162.5028 |  |  | 2 | 40651555 | 69.2565 | 102.6%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 39700209 | 67.6357 | 100.2%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100%
[6 empty] | RSD: | 1.419%
[6 empty] | n: | 6
Comment: | The content of | Amodiaquine |  | in the sample as a percentage of the stated label claim is |  | 100%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Amodiaquine HCl
Code: | A7 1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 355.79 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 428.79 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.82975 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 18042489 | 27876745.3128 | 17014697 | 28010714.1546
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 18126127 | 28005971.1211 | 17104638 | 28158780.949 |  | 0.5173%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 18032777 | 27861739.6808 | 17030396 | 28036558.8818
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 18067131 | 27914818.7049 | 17049910.3333 | 28068684.6618
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 12.2 |  | 11.45
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 10.123 |  | 9.5007
Standard Dilution Factor | 208.3333 | Purity correction: | 10.1129 |  | 9.4912
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.048542 |  | 0.045558
|  | Desired Concentration (mg/mL): | 0.075
|  | Desired Weight as free base (mg): | 15.625
|  | Desired Weight as salt (mg): | 18.8309
|  | Average Normalised Peak Area: | 27991751.6833
|  | RSD: | 0.3933%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 23404393 | 56.4379 | 83.61%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 23543576 | 56.7736 | 84.11%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 23489602 | 56.6434 | 83.92%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 23383537 | 56.3876 | 83.54%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 23768908 | 57.3169 | 84.91%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 23533788 | 56.75 | 84.07%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 84.03%
Sample Dilution Factor | 900 |  |  | RSD: | 0.5879%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Amodiaquine |  | dissolved as a percentage of the stated label claim is |  | 84.03%
| Name |  |  | Date |  | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:
0

## Artesunate
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Artesunate & Amodiaquine Tablets
Laboratory Ref No: | NDQD201508213 | 21
Active Ingredient: | Artesunate
Label Claim: | Each bilayered tablet contains Artesunate 25 mg & Amodiaquine HCl 67.5mg
Date Analysis Started: | 2015-11-04
Date Analysis Completed: | 2015-11-30
Analysis Data
Reference Substance: | Artesunate
Code: | A15 2
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 3376016 | 2969270.4829 | 3669490 | 2953922.8637
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 3387181 | 2979090.3134 | 3650784 | 2938864.6183 |  | 1.225%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 3388021 | 2979829.1094 | 3646394 | 2935330.6881
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 3383739.3333 | 2976063.3019 | 3655556 | 2942706.0567
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 22.9 |  | 25.02
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 22.9 |  | 25.02
Standard Dilution Factor | 5 | Purity correction (mg): | 22.7397 |  | 24.8449
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 4.5479 |  | 4.969
|  | Desired Concentration (mg/mL): | 4
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 2959384.6793
|  | RSD: | 0.6645%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each bilayered tablet contains Artesunate 25 mg & Amodiaquine HCl 67.5mg
Each Tablet contains | 25 | Artesunate
Average Tablet Content Weight (mg): | 175.503
Initial Sample dilution Volume (mL): | 5 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 178.24 | 1 | 3914036 | 26.0455 | 104.2%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3914457 | 26.0483 | 104.2%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3929448 | 26.1481 | 104.6%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 175.87 | 1 | 3640826 | 24.5539 | 98.22%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3663429 | 24.7064 | 98.83%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3683988 | 24.845 | 99.38%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5 | Assay Smp C | 175.79 | 1 | 3729887 | 25.166 | 100.7%
Desired Sample Weight (mg): | 140.4024 |  |  | 2 | 3741135 | 25.2419 | 101%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 3750607 | 25.3058 | 101.2%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100.5%
[6 empty] | RSD: | 1.974%
[6 empty] | n: | 7
Comment: | The content of | Artesunate |  | in the sample as a percentage of the stated label claim is |  | 100.5%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Artesunate
Code: | A15 2
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 1039155 | 1142445.8106 | 1102951 | 1109838.212
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 1040823 | 1144279.6079 | 1119345 | 1126334.5819 |  | 2.514%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 1048484 | 1152702.1025 | 1117130 | 1124105.7506
3rd Dilution Standard Volume (mL: | 1 | 4 |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 1042820.6667 | 1146475.8403 | 1113142 | 1120092.8482
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 22.9 |  | 25.02
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 22.9 |  | 25.02
Standard Dilution Factor | 500 | Purity correction: | 22.7397 |  | 24.8449
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.045479 |  | 0.04969
|  | Desired Concentration (mg/mL): | 0.05
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Peak Area: | 1133284.3443
|  | RSD: | 1.403%
|  | n: | 6
Medium Volume (mL): | 500 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 1201469 | 26.5041 | 106%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 1175523 | 25.9318 | 103.7%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 1169485 | 25.7986 | 103.2%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 1179741 | 26.0248 | 104.1%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 1202099 | 26.518 | 106.1%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 1174175 | 25.902 | 103.6%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 104.5%
Sample Dilution Factor | 500 |  |  | RSD: | 1.212%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Artesunate |  | dissolved as a percentage of the stated label claim is |  | 104.5%
| Name |  |  | Date |  | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:
0