## SST (AH)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Arteunate & Amodiaquine Tablets
Reference Substance: | Amodiaquine HCl
% age Purity: | 99.9
Weight (mg): | 11.8
Standard Conc (mg/mL): | 0.1888
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 28822227 | 2537.5 | 2 | 2
2 | 28986515 | 2549.7 | 1.9 | 2
3 | 28910724 | 2573.7 | 1.9 | 2
4 | 28919077 | 2590.1 | 1.9 | 2
5 | 28943023 | 2576.2 | 1.9 | 2
6 | 28875805 | 2574.1 | 1.9 | 2
Averages: | 28909561.8333 | 2566.8833 | 1.9167 | 2
RSD: | 0.1951%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Amodiaquine HCl
% age Purity: | 99.9
Weight (mg): | 11.8
Standard Conc (mg/mL): | 0.05664
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 28822227 | 2537.5 | 2 | 2
2 | 28986515 | 2549.7 | 1.9 | 2
3 | 28910724 | 2573.7 | 1.9 | 2
4 | 28919077 | 2590.1 | 1.9 | 2
5 | 28943023 | 2576.2 | 1.9 | 2
6 | 28875805 | 2574.1 | 1.9 | 2
Averages: | 28909561.8333 | 2566.8833 | 1.9167 | 2
RSD: | 0.1951%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:

## SST (A)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Arteunate & Amodiaquine Tablets
Reference Substance: | Artesunate
% age Purity: | 99.3
Weight (mg): | 21.88
Standard Conc (mg/mL): | 4.376
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3175409 | 8053.2 | 1 | 8.5
2 | 3172253 | 8024.7 | 1 | 8.5
3 | 3177874 | 8058.9 | 1 | 8.5
4 | 3179284 | 8077.6 | 1 | 8.5
5 | 3178130 | 8061.1 | 1 | 8.5
6 | 3183344 | 8100.5 | 1 | 8.5
Averages: | 3177715.6667 | 8062.6667 | 1 | 8.5
RSD: | 0.1173%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Artesunate
% age Purity: | 99.3
Weight (mg): | 21.88
Standard Conc (mg/mL): | 0.04376
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3175409 | 8053.2 | 1 | 8.5
2 | 3172253 | 8024.7 | 1 | 8.5
3 | 3177874 | 8058.9 | 1 | 8.5
4 | 3179284 | 8077.6 | 1 | 8.5
5 | 3178130 | 8061.1 | 1 | 8.5
6 | 3183344 | 8100.5 | 1 | 8.5
Averages: | 3177715.6667 | 8062.6667 | 1 | 8.5
RSD: | 0.1173%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | Arteunate & Amodiaquine Tablets
Laboratory Ref No: |  | NDQD201508216
Active Ingredient: |  | Arteunate & Amodiaquine
Label Claim: |  | Each bilayered tablet contains Artesunate 100mg and Amodiaquine 270mg
Date Analysis Started: |  | 2015-11-04
Date Analysis Completed: |  | 2015-11-30
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 712.1 | 0.9294%
|  | 707.58 | 0.2887%
|  | 697.36 | -1.16%
|  | 705.69 | 0.02084%
|  | 703.61 | -0.274%
|  | 711.86 | 0.8953%
|  | 701.07 | -0.634%
|  | 712.09 | 0.9279%
|  | 698.4 | -1.012%
|  | 718.52 | 1.839%
|  | 705.94 | 0.05627%
|  | 702.8 | -0.3888%
|  | 706.81 | 0.1796%
|  | 714.15 | 1.22%
|  | 709.7 | 0.5892%
|  | 700.51 | -0.7134%
|  | 709.36 | 0.541%
|  | 693.97 | -1.64%
|  | 695.42 | -1.435%
|  | 703.92 | -0.23%
| Total | 14110.86
| Average | 705.543
| Average | % Deviation from mean
| 705.543 | -5% | 670.2658
|  | 5% | 740.8201
| Name |  | Date | Signature
Analysed by: | Bugigi |  | 2015-12-01
Reviewed By:

## Amodiaquine HCl
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Artesunate & Amodiaquine Tablets
Laboratory Ref No: | NDQD201508216 | 21
Active Ingredient: | Amodiaquine
Label Claim: | Each bilayered tablet contains Artesunate 100mg and Amodiaquine 270mg
Date Analysis Started: | 2015-11-04
Date Analysis Completed: | 2015-11-30
Analysis Data
Reference Substance: | Amodiaquine HCl
Code: | A7 1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 355.79 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 428.79 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.82975 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 28263753 | 36119627.0643 | 31331416 | 36011534.1716
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 28348561 | 36228007.3396 | 31286024 | 35959361.7591 |  | 0.6093%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 28310763 | 36179703.4338 | 31288065 | 35961707.6327
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 28307692.3333 | 36175779.2792 | 31301835 | 35977534.5211
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 11.8 |  | 13.12
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 9.7911 |  | 10.8864
Standard Dilution Factor | 62.5 | Purity correction (mg): | 9.7813 |  | 10.8755
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.1565 |  | 0.17401
|  | Desired Concentration (mg/mL): | 0.2
|  | Desired Weight as free base (mg): | 12.5
|  | Desired Weight as salt (mg): | 15.0647
|  | Average Normalised Response: | 36076656.9002
|  | RSD: | 0.3199%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each bilayered tablet contains Artesunate 100mg and Amodiaquine 270mg
Each Tablet contains | 270 | Amodiaquine
Average Tablet Content Weight (mg): | 705.543
Initial Sample dilution Volume (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 4 | Assay Smp A | 706.17 | 1 | 39160970 | 271.1324 | 100.4%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 39761947 | 275.2933 | 102%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 39751473 | 275.2207 | 101.9%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 704.13 | 1 | 38993986 | 270.7584 | 100.3%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 39088114 | 271.412 | 100.5%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 39280768 | 272.7497 | 101%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1250 | Assay Smp C | 700.86 | 1 | 39113415 | 272.8548 | 101.1%
Desired Sample Weight (mg): | 653.2806 |  |  | 2 | 38993049 | 272.0152 | 100.7%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 39032401 | 272.2897 | 100.8%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 101%
[6 empty] | RSD: | 0.6026%
[6 empty] | n: | 9
Comment: | The content of | Amodiaquine |  | in the sample as a percentage of the stated label claim is |  | 101%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Amodiaquine HCl
Code: | A7 1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 355.79 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 428.79 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.82975 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 17041876 | 27223304.5978 | 18825247 | 27046576.2555
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 16927886 | 27041212.879 | 18918178 | 27180091.9208 |  | 0.1834%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 17078749 | 27282206.8519 | 18921968 | 27185537.083
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 17016170.3333 | 27182241.4429 | 18888464.3333 | 27137401.7531
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 11.8 |  | 13.12
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 9.7911 |  | 10.8864
Standard Dilution Factor | 208.3333 | Purity correction: | 9.7813 |  | 10.8755
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.04695 |  | 0.052202
|  | Desired Concentration (mg/mL): | 0.075
|  | Desired Weight as free base (mg): | 15.625
|  | Desired Weight as salt (mg): | 18.8309
|  | Average Normalised Peak Area: | 27159821.598
|  | RSD: | 0.3569%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 5 | 1 | 23341122 | 232.0377 | 85.94%
1st Dilution Sample Final Volume (mL): | 20 | 2 | 23334193 | 231.9688 | 85.91%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 23321431 | 231.842 | 85.87%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 23263206 | 231.2631 | 85.65%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 23335115 | 231.978 | 85.92%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 23249942 | 231.1313 | 85.6%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 85.82%
Sample Dilution Factor | 3600 |  |  | RSD: | 0.1724%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Amodiaquine |  | dissolved as a percentage of the stated label claim is |  | 85.82%
| Name |  |  | Date |  | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:
0

## Artesunate
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Artesunate & Amodiaquine Tablets
Laboratory Ref No: | NDQD201508216 | 21
Active Ingredient: | Artesunate
Label Claim: | Each bilayered tablet contains Artesunate 100mg and Amodiaquine 270mg
Date Analysis Started: | 2015-11-04
Date Analysis Completed: | 2015-11-30
Analysis Data
Reference Substance: | Artesunate
Code: | A15 2
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 3182882 | 2929907.8927 | 3653857 | 2934301.6747
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 3186011 | 2932788.2011 | 3670313 | 2947516.9889 |  | 0.2948%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 3193921 | 2940069.5177 | 3665418 | 2943585.963
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 3187604.6667 | 2934255.2039 | 3663196 | 2941801.5422
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 21.88 |  | 25.08
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 21.88 |  | 25.08
Standard Dilution Factor | 5 | Purity correction (mg): | 21.7268 |  | 24.9044
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 4.3454 |  | 4.9809
|  | Desired Concentration (mg/mL): | 4
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 2938028.373
|  | RSD: | 0.232%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each bilayered tablet contains Artesunate 100mg and Amodiaquine 270mg
Each Tablet contains | 100 | Artesunate
Average Tablet Content Weight (mg): | 705.543
Initial Sample dilution Volume (mL): | 20 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 693.49 | 1 | 3486104 | 96.5734 | 96.57%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3489703 | 96.6731 | 96.67%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3486947 | 96.5968 | 96.6%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 709.54 | 1 | 3602366 | 97.5368 | 97.54%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3604546 | 97.5958 | 97.6%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3620448 | 98.0264 | 98.03%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 20 | Assay Smp C | 718.96 | 1 | 3719401 | 99.3861 | 99.39%
Desired Sample Weight (mg): | 564.4344 |  |  | 2 | 3727205 | 99.5947 | 99.59%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 3727645 | 99.6064 | 99.61%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 97.95%
[6 empty] | RSD: | 1.31%
[6 empty] | n: | 9
Comment: | The content of | Artesunate |  | in the sample as a percentage of the stated label claim is |  | 97.95%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Artesunate
Code: | A15 2
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 971735 | 1118127.3945 | 1126249 | 1130570.4927
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 961876 | 1106783.1309 | 1112255 | 1116522.7967 |  | 0.8601%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 970867 | 1117128.6298 | 1115743 | 1120024.1804
3rd Dilution Standard Volume (mL: | 1 | 4 |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 968159.3333 | 1114013.0518 | 1118082.3333 | 1122372.4899
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 21.88 |  | 25.08
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 21.88 |  | 25.08
Standard Dilution Factor | 500 | Purity correction: | 21.7268 |  | 24.9044
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.043454 |  | 0.049809
|  | Desired Concentration (mg/mL): | 0.05
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Peak Area: | 1118192.7708
|  | RSD: | 0.6819%
|  | n: | 6
Medium Volume (mL): | 500 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 5 | 1 | 1100717 | 98.4371 | 98.44%
1st Dilution Sample Final Volume (mL): | 20 | 2 | 1108828 | 99.1625 | 99.16%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 1084352 | 96.9736 | 96.97%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 1097338 | 98.135 | 98.13%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 1077861 | 96.3931 | 96.39%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 1095082 | 97.9332 | 97.93%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 97.84%
Sample Dilution Factor | 2000 |  |  | RSD: | 1.027%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Artesunate |  | dissolved as a percentage of the stated label claim is |  | 97.84%
| Name |  |  | Date |  | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:
0