## SST (AH)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Arteunate & Amodiaquine Tablets
Reference Substance: | Amodiaquine HCl
% age Purity: | 99.9
Weight (mg): | 11.8
Standard Conc (mg/mL): | 0.1888
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 28822227 | 2537.5 | 2 | 2
2 | 28986515 | 2549.7 | 1.9 | 2
3 | 28910724 | 2573.7 | 1.9 | 2
4 | 28919077 | 2590.1 | 1.9 | 2
5 | 28943023 | 2576.2 | 1.9 | 2
6 | 28875805 | 2574.1 | 1.9 | 2
Averages: | 28909561.8333 | 2566.8833 | 1.9167 | 2
RSD: | 0.1951%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Amodiaquine HCl
% age Purity: | 99.9
Weight (mg): | 11.8
Standard Conc (mg/mL): | 0.05664
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 28822227 | 2537.5 | 2 | 2
2 | 28986515 | 2549.7 | 1.9 | 2
3 | 28910724 | 2573.7 | 1.9 | 2
4 | 28919077 | 2590.1 | 1.9 | 2
5 | 28943023 | 2576.2 | 1.9 | 2
6 | 28875805 | 2574.1 | 1.9 | 2
Averages: | 28909561.8333 | 2566.8833 | 1.9167 | 2
RSD: | 0.1951%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:

## SST (A)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Arteunate & Amodiaquine Tablets
Reference Substance: | Artesunate
% age Purity: | 99.3
Weight (mg): | 21.88
Standard Conc (mg/mL): | 4.376
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3175409 | 8053.2 | 1 | 8.5
2 | 3172253 | 8024.7 | 1 | 8.5
3 | 3177874 | 8058.9 | 1 | 8.5
4 | 3179284 | 8077.6 | 1 | 8.5
5 | 3178130 | 8061.1 | 1 | 8.5
6 | 3183344 | 8100.5 | 1 | 8.5
Averages: | 3177715.6667 | 8062.6667 | 1 | 8.5
RSD: | 0.1173%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Artesunate
% age Purity: | 99.3
Weight (mg): | 21.88
Standard Conc (mg/mL): | 0.04376
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3175409 | 8053.2 | 1 | 8.5
2 | 3172253 | 8024.7 | 1 | 8.5
3 | 3177874 | 8058.9 | 1 | 8.5
4 | 3179284 | 8077.6 | 1 | 8.5
5 | 3178130 | 8061.1 | 1 | 8.5
6 | 3183344 | 8100.5 | 1 | 8.5
Averages: | 3177715.6667 | 8062.6667 | 1 | 8.5
RSD: | 0.1173%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | Arteunate & Amodiaquine Tablets
Laboratory Ref No: |  | NDQD201508219
Active Ingredient: |  | Arteunate & Amodiaquine
Label Claim: |  | Each bilayered tablet contains Artesunate 50mg and Amodiaquine 135mg
Date Analysis Started: |  | 2015-11-04
Date Analysis Completed: |  | 2015-11-30
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 369.08 | 2.225%
|  | 362.06 | 0.2803%
|  | 362.16 | 0.308%
|  | 365.75 | 1.302%
|  | 361.49 | 0.1224%
|  | 363.5 | 0.6791%
|  | 361.72 | 0.1861%
|  | 362.72 | 0.4631%
|  | 358.22 | -0.7833%
|  | 355.47 | -1.545%
|  | 356.36 | -1.298%
|  | 361.57 | 0.1446%
|  | 359.44 | -0.4454%
|  | 363.18 | 0.5905%
|  | 361.5 | 0.1252%
|  | 356.1 | -1.37%
|  | 357.92 | -0.8664%
|  | 358.63 | -0.6697%
|  | 367.7 | 1.842%
|  | 356.39 | -1.29%
| Total | 7220.96
| Average | 361.048
| Average | % Deviation from mean
| 361.048 | -5% | 342.9956
|  | 5% | 379.1004
| Name |  | Date | Signature
Analysed by: | Bugigi |  | 2015-12-01
Reviewed By:

## Amodiaquine HCl
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Artesunate & Amodiaquine Tablets
Laboratory Ref No: | NDQD201508219 | 21
Active Ingredient: | Amodiaquine
Label Claim: | Each bilayered tablet contains Artesunate 50mg and Amodiaquine 135mg
Date Analysis Started: | 2015-11-04
Date Analysis Completed: | 2015-11-30
Analysis Data
Reference Substance: | Amodiaquine HCl
Code: | A7 1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 355.79 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 428.79 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.82975 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 28263753 | 36119627.0643 | 31331416 | 36011534.1716
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 28348561 | 36228007.3396 | 31286024 | 35959361.7591 |  | 0.6093%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 28310763 | 36179703.4338 | 31288065 | 35961707.6327
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 28307692.3333 | 36175779.2792 | 31301835 | 35977534.5211
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 11.8 |  | 13.12
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 9.7911 |  | 10.8864
Standard Dilution Factor | 62.5 | Purity correction (mg): | 9.7813 |  | 10.8755
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.1565 |  | 0.17401
|  | Desired Concentration (mg/mL): | 0.2
|  | Desired Weight as free base (mg): | 12.5
|  | Desired Weight as salt (mg): | 15.0647
|  | Average Normalised Response: | 36076656.9002
|  | RSD: | 0.3199%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each bilayered tablet contains Artesunate 50mg and Amodiaquine 135mg
Each Tablet contains | 135 | Amodiaquine
Average Tablet Content Weight (mg): | 361.048
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 3 | Assay Smp A | 356.56 | 1 | 34538610 | 129.2557 | 95.74%
1st Dilution Sample Final Volume (mL): | 20 |  |  | 2 | 34395919 | 128.7217 | 95.35%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 34662906 | 129.7208 | 96.09%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 362.27 | 1 | 34227020 | 126.0707 | 93.39%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 34406695 | 126.7325 | 93.88%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 34547231 | 127.2501 | 94.26%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 666.6667 | Assay Smp C | 355.42 | 1 | 34305317 | 128.7944 | 95.4%
Desired Sample Weight (mg): | 356.5906 |  |  | 2 | 34283126 | 128.7111 | 95.34%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 34236257 | 128.5351 | 95.21%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 94.96%
[6 empty] | RSD: | 0.956%
[6 empty] | n: | 9
Comment: | The content of | Amodiaquine |  | in the sample as a percentage of the stated label claim is |  | 94.96%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Amodiaquine HCl
Code: | A7 1
% age Purity: | 99.9
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.9
Mwt of compound in free base form: | 355.79 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 428.79 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.82975 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 17041876 | 27223304.5978 | 18825247 | 27046576.2555
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 16927886 | 27041212.879 | 18918178 | 27180091.9208 |  | 0.1834%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 17078749 | 27282206.8519 | 18921968 | 27185537.083
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 17016170.3333 | 27182241.4429 | 18888464.3333 | 27137401.7531
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 11.8 |  | 13.12
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 9.7911 |  | 10.8864
Standard Dilution Factor | 208.3333 | Purity correction: | 9.7813 | 1126249 | 10.8755
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.04695 | 1112255 | 0.052202
|  | Desired Concentration (mg/mL): | 0.075 | 1115743
|  | Desired Weight as free base (mg): | 15.625
|  | Desired Weight as salt (mg): | 18.8309
|  | Average Normalised Peak Area: | 27159821.598
|  | RSD: | 0.3569%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 22112544 | 109.9121 | 81.42%
1st Dilution Sample Final Volume (mL): | 20 | 2 | 21420932 | 106.4744 | 78.87%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 21373979 | 106.241 | 78.7%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 21289315 | 105.8202 | 78.39%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 21530899 | 107.021 | 79.27%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 21824627 | 108.481 | 80.36%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 79.5%
Sample Dilution Factor | 1800 |  |  | RSD: | 1.46%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Amodiaquine |  | dissolved as a percentage of the stated label claim is |  | 79.5%
| Name |  |  | Date |  | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:
0

## Artesunate
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Artesunate & Amodiaquine Tablets
Laboratory Ref No: | NDQD201508219 | 21
Active Ingredient: | Artesunate
Label Claim: | Each bilayered tablet contains Artesunate 50mg and Amodiaquine 135mg
Date Analysis Started: | 2015-11-04
Date Analysis Completed: | 2015-11-30
Analysis Data
Reference Substance: | Artesunate
Code: | A15 2
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 | 3182882 | 2929907.8927 | 3653857 | 2934301.6747
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 3186011 | 2932788.2011 | 3670313 | 2947516.9889 |  | 0.2948%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 3193921 | 2940069.5177 | 3665418 | 2943585.963
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 3187604.6667 | 2934255.2039 | 3663196 | 2941801.5422
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 21.88 |  | 25.08
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 21.88 |  | 25.08
Standard Dilution Factor | 5 | Purity correction (mg): | 21.7268 |  | 24.9044
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 4.3454 |  | 4.9809
|  | Desired Concentration (mg/mL): | 4
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 2938028.373
|  | RSD: | 0.232%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each bilayered tablet contains Artesunate 50mg and Amodiaquine 135mg
Each Tablet contains | 50 | Artesunate
Average Tablet Content Weight (mg): | 361.048
Initial Sample dilution Volume (mL): | 10 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 363.32 | 1 | 3740345 | 50.6048 | 101.2%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3751870 | 50.7607 | 101.5%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3758420 | 50.8493 | 101.7%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 362.47 | 1 | 3653867 | 49.5507 | 99.1%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 3668650 | 49.7512 | 99.5%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 3680829 | 49.9163 | 99.83%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 10 | Assay Smp C | 368.73 | 1 | 3793981 | 50.5773 | 101.2%
Desired Sample Weight (mg): | 288.8384 |  |  | 2 | 3808197 | 50.7668 | 101.5%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 3812174 | 50.8198 | 101.6%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100.8%
[6 empty] | RSD: | 1.015%
[6 empty] | n: | 9
Comment: | The content of | Artesunate |  | in the sample as a percentage of the stated label claim is |  | 100.8%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Artesunate
Code: | A15 2
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 971735 | 1118127.3945 | 1126249 | 1130570.4927
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 961876 | 1106783.1309 | 1112255 | 1116522.7967 |  | 0.8601%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 970867 | 1117128.6298 | 1115743 | 1120024.1804
3rd Dilution Standard Volume (mL: | 1 | 4 |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 968159.3333 | 1114013.0518 | 1118082.3333 | 1122372.4899
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 21.88 |  | 25.08
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 21.88 |  | 25.08
Standard Dilution Factor | 500 | Purity correction: | 21.7268 |  | 24.9044
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.043454 |  | 0.049809
|  | Desired Concentration (mg/mL): | 0.05
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Peak Area: | 1118192.7708
|  | RSD: | 0.6819%
|  | n: | 6
Medium Volume (mL): | 500 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 1153529 | 51.5801 | 103.2%
1st Dilution Sample Final Volume (mL): | 20 | 2 | 1164663 | 52.0779 | 104.2%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 1156625 | 51.7185 | 103.4%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 1160709 | 51.9011 | 103.8%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 1161029 | 51.9154 | 103.8%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 1157799 | 51.771 | 103.5%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 103.7%
Sample Dilution Factor | 1000 |  |  | RSD: | 0.3366%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Comment: | The amount of | Artesunate |  | dissolved as a percentage of the stated label claim is |  | 103.7%
| Name |  |  | Date |  | Signature
Analysed by: | Bugigi |  |  | 2015-12-01
Reviewed By:
0