## SST-S1 Dissolution
HPLC System Suitability Report
Analysis Data | Dissolution
Sample(s) | ELAXIC P TABLETS
Reference Substance: | Paracetamol
% age Purity: | 99.84
Weight (mg): | 23.55
Standard Conc (mg/mL): | 0.0471
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 14575001 | 8533.09 | 1.14 | 2.83
2 | 14519963 | 8556.69 | 1.16 | 2.83
3 | 14535446 | 8551.14 | 1.18 | 2.83
4 | 14532918 | 8540.91 | 1.15 | 2.83
5 | 14539694 | 8508.8 | 1.21 | 2.83
6 | 14518508 | 8575.65 | 1.18 | 2.83
Averages: | 4480164.3333 | 8309.1633 | 1.3133 | 6.39
RSD: | 0.4577%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Aceclofenac
% age Purity: | 99.44
Weight (mg): | 9.09
Standard Conc (mg/mL): | 0.00909
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 5607418 | 10241.34 | 1.06 | 6.93
2 | 5594365 | 10227.36 | 1.04 | 6.93
3 | 5611253 | 10217.77 | 1.04 | 6.93
4 | 5614131 | 10233.36 | 1.05 | 6.93
5 | 5612671 | 10246.59 | 1.04 | 6.93
6 | 5621054 | 10241.96 | 1.04 | 6.94
Averages: | 4480164.3333 | 8309.1633 | 1.3133 | 6.39
RSD: | 0.1994%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO/LORNA |  |  | 2016-08-08
Reviewed By:

## SST-S2 Dissolution
HPLC System Suitability Report
Analysis Data | Dissolution
Sample(s) | ELAXIC P TABLETS
Reference Substance: | Paracetamol
% age Purity: | 99.84
Weight (mg): | 12.39
Standard Conc (mg/mL): | 0.04956
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 15484826 | 7735.1 | 1.1 | 2.84
2 | 15475869 | 7782.52 | 1.1 | 2.84
3 | 15474599 | 7762.21 | 1.08 | 2.84
4 | 15462310 | 7730.17 | 1.13 | 2.84
5 | 15436330 | 7780.56 | 1.08 | 2.84
6 | 15459183 | 7742.08 | 1.12 | 2.84
Averages: | 15465519.5 | 7755.44 | 1.1017 | 2.84
RSD: | 0.1107%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Aceclofenac
% age Purity: | 99.44
Weight (mg): | 11.25
Standard Conc (mg/mL): | 0.01125
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 7105206 | 8274.22 | 0.93 | 7
2 | 7105768 | 8301.73 | 0.93 | 7.01
3 | 7112525 | 8309.83 | 0.92 | 7.01
4 | 7106058 | 8311.21 | 0.93 | 7
5 | 7101111 | 8320.25 | 0.93 | 7
6 | 7113746 | 8326.83 | 0.94 | 7
Averages: | 7107402.3333 | 8307.345 | 0.93 | 7.0033
RSD: | 0.06758%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO/JOYFRIDA |  |  | 2016-01-08
Reviewed By:

## SST-Aceclofenac Assay
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | ELAXIC P TABLETS
Reference Substance: | Aceclofenac
% age Purity: | 99.44
Weight (mg): | 27.69
Standard Conc (mg/mL): | 0.0033228
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 2070115 | 10331.27 | 1.01 | 6.95
2 | 2067891 | 10361.88 | 1.03 | 6.95
3 | 2073498 | 10322.54 | 1.01 | 6.95
4 | 2074066 | 10312.31 | 1.02 | 6.95
5 | 2072143 | 10316.98 | 1.01 | 6.95
6 | 2072147 | 10317.4 | 1.01 | 6.95
Averages: | 2071643.3333 | 10327.0633 | 1.015 | 6.95
RSD: | 0.1106%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO |  |  | 2016-02-08
Reviewed By:

## SST-Paracetamol Assay
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | ELAXIC P TABLETS
Reference Substance: | Paracetamol
% age Purity: | 99.84
Weight (mg): | 23.03
Standard Conc (mg/mL): | 0.013818
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4364827 | 8616.38 | 1.17 | 2.83
2 | 4366340 | 8575.66 | 1.14 | 2.84
3 | 4363457 | 8615.89 | 1.16 | 2.84
4 | 4375466 | 8584.49 | 1.17 | 2.83
5 | 4367620 | 8613.45 | 1.17 | 2.84
6 | 4368738 | 8611.02 | 1.16 | 2.84
Averages: | 4367741.3333 | 8602.815 | 1.1617 | 2.8367
RSD: | 0.09686%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO |  |  | 2016-02-08
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | ELAXIC P TABLETS
Laboratory Ref No: |  | NDQD201510452
Active Ingredient: |  | Aceclofenac BP, Paracetamol BP
Label Claim: |  | Each Tablet contains: Aceclofenac, Paracetamol
Date Analysis Started: |  | 2015-10-26 12:27:02
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 718.17 | 0.1715%
|  | 704.09 | -1.792%
|  | 730.34 | 1.869%
|  | 720.78 | 0.5355%
|  | 702.89 | -1.96%
|  | 705.33 | -1.619%
|  | 720.93 | 0.5565%
|  | 720.05 | 0.4337%
|  | 724.48 | 1.052%
|  | 725.21 | 1.153%
|  | 700.64 | -2.274%
|  | 727.84 | 1.52%
|  | 726.71 | 1.363%
|  | 724.38 | 1.038%
|  | 716.47 | -0.06563%
|  | 717.14 | 0.02783%
|  | 725.61 | 1.209%
|  | 709.54 | -1.032%
|  | 708.05 | -1.24%
|  | 710.16 | -0.9458%
| Total | 14338.81
| Average | 716.9405
| Average | % Deviation from mean
| 716.9405 | -5% | 681.0935
|  | 5% | 752.7875
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Aceclofenac
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ELAXIC P TABLETS
Laboratory Ref No: | NDQD201510452 | 11
Active Ingredient: | Aceclofenac
Label Claim: | Each Tablet contains: Aceclofenac BP 100mg , Paracetamol BP 500mg
Date Analysis Started: | 2016-08-01
Date Analysis Completed: | 2016-08-02
Analysis Data
Reference Substance: | Aceclofenac
Code: | A52-7
% age Purity: | 99.44
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.44
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 20 | 1 | 2085577 | 1893573.4951 | 1870794 | 1837953.7367
2nd Dilution Standard Transfer Volume (mL): | 3 | 2 | 2076265 | 1885118.7815 | 1868022 | 1835230.3969
2nd Dilution Standard Final Volume (mL): | 100 | 3 | 2075646 | 1884556.7682 | 1867770 | 1834982.8206
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 2079162.6667 | 1887749.6816 | 1868862 | 1836055.6514
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 27.69 |  | 25.59
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 27.69 |  | 25.59
Standard Dilution Factor | 8333.3333 | Purity correction (mg): | 27.5349 |  | 25.4467
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.0033042 |  | 0.0030536
|  | Desired Concetration (mg/mL): | 0.003
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Response: | 1861902.6665
|  | RSD: | 1.531%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each Tablet contains: Aceclofenac BP 100mg , Paracetamol BP 500mg
Each Tablet contains | 100 | Aceclofenac
Average Tablet Content Weight (mg): | 716.9405
Initial Sample dilution (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 708.68 | 1 | 2232855 | 90.9909 | 90.99%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 2230748 | 90.905 | 90.9%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 2225894 | 90.7072 | 90.71%
2nd Dilution Sample Final Volume (mL): | 10 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 713.66 | 1 | 2292418 | 92.7662 | 92.77%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 2285841 | 92.5001 | 92.5%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 2286437 | 92.5242 | 92.52%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 25000 | Assay Smp C | 716.7 | 1 |  | - | -
Desired Sample Weight (mg) | 537.7054 |  |  | 2 |  | - | -
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 |  | - | -
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 91.73%
[6 empty] | RSD: | 1.042%
[6 empty] | n: | 6
Comment | The content of | Aceclofenac |  | in the sample as a percentage of the stated label claim is |  | 91.73%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Aceclofenac
Code: | A52-7
% age Purity: | 99.44
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.44
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 5575543 | 6853614.3192 | 7693171 | 7022962.3948
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 5579615 | 6858619.7361 | 7689868 | 7019947.1434
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 5587799 | 6868679.7392 | 7710678 | 7038944.2315
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 5580985.6667 | 6860304.5981 | 7697905.6667 | 7027284.5899
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 9.09 |  | 12.24
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 9.09 |  | 12.24
Standard Dilution Factor | 1000 | Purity correction (mg): | 9.0391 |  | 12.1715
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.0090391 |  | 0.012171
|  | Desired Concetration (mg/mL): | 0.011111
|  | Desired Weight as free base (mg): | 11.1111
|  | Desired Weight as salt (mg): | 11.1111
|  | Average Normalised Peak Area: | 6943794.594
|  | RSD: | 1.322%
|  | n: | 6
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 5 | 1 | 3869625 | 55.7278 | 55.73%
1st Dilution Sample Final Volume (mL): | 50 | 2 | 4232236 | 60.9499 | 60.95%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 4551081 | 65.5417 | 65.54%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4243824 | 61.1168 | 61.12%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 4046613 | 58.2767 | 58.28%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 4377879 | 63.0474 | 63.05%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 60.78%
Sample Dilution Factor | 9000 |  |  | RSD: | 5.688%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Aceclofenac
Code: | A52-7
% age Purity: | 99.44
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.44
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 7090586 | 7042485.8218 | 6939826 | 7004825.7394
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 7110414 | 7062179.3153 | 6944067 | 7009106.4614
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 7113856 | 7065597.9659 | 6930620 | 6995533.5142
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 7104952 | 7056754.3677 | 6938171 | 7003155.2383
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 11.25 |  | 11.07
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 11.25 |  | 11.07
Standard Dilution Factor | 1000 | Purity correction: | 11.187 |  | 11.008
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.011187 |  | 0.011008
|  | Desired Concetration (mg/mL): | 0.011111
|  | Desired Weight as free base (mg): | 11.1111
|  | Desired Weight as salt (mg): | 11.1111
|  | Average Normalised Peak Area: | 7029954.803
|  | RSD: | 0.4369%
|  | n: | 6
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 4294684 | 61.0912 | 61.09%
1st Dilution Sample Final Volume (mL): | 100 | 2 | 4605995 | 65.5196 | 65.52%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 4579140 | 65.1375 | 65.14%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4620404 | 65.7245 | 65.72%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 4214394 | 59.9491 | 59.95%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 4481819 | 63.7532 | 63.75%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 63.53%
Sample Dilution Factor | 9000 |  |  | RSD: | 3.867%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 62.15%
RSD: | 5.149%
n: | 12
Comment | The amount of | Aceclofenac |  | dissolved as a percentage of the stated label claim is |  | 62.15%
| Name |  |  | Date |  | Signature
Analysed by: |  | RUTTO/JOYFRIDA
Reviewed By:
5

## Paracetamol
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ELAXIC P TABLETS
Laboratory Ref No: | NDQD201510452 | 11
Active Ingredient: | Paracetamol
Label Claim: | Each Tablet contains: Aceclofenac BP 100mg , Paracetamol BP 500mg
Date Analysis Started: | 2016-08-01
Date Analysis Completed: | 2016-08-02
Analysis Data
Reference Substance: | Paracetamol
Code: | P1-3
% age Purity: | 99.84
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.84
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4379757 | 4762023.2755 | 4845072 | 4680590.7971
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4365685 | 4746723.0678 | 4848216 | 4683628.0642
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4365071 | 4746055.4777 | 4839499 | 4675206.9901
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4370171 | 4751600.607 | 4844262.3333 | 4679808.6171
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 23.03 |  | 25.92
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 23.03 |  | 25.92
Standard Dilution Factor | 1666.6667 | Purity correction (mg): | 22.9932 |  | 25.8785
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.013796 |  | 0.015527
|  | Desired Concetration (mg/mL): | 0.015
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Response: | 4715704.6121
|  | RSD: | 0.8445%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each Tablet contains: Aceclofenac BP 100mg , Paracetamol BP 500mg
Each Tablet contains | 500 | Paracetamol
Average Tablet Content Weight (mg): | 716.9405
\
Initial Sample dilution (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 4 | Assay Smp A | 708.68 | 1 | 6149111 | 494.6864 | 98.94%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 6147979 | 494.5953 | 98.92%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 6145771 | 494.4177 | 98.88%
2nd Dilution Sample Final Volume (mL): | 10 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 713.66 | 1 | 6311646 | 504.2188 | 100.8%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 6301149 | 503.3803 | 100.7%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 6304060 | 503.6128 | 100.7%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 25000 | Assay Smp C | 716.7 | 1 | 6288787 | 500.2617 | 100.1%
Desired Sample Weight (mg) | 537.7054 |  |  | 2 | 6278805 | 499.4677 | 99.89%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 6280652 | 499.6146 | 99.92%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 99.87%
[6 empty] | RSD: | 0.8001%
[6 empty] | n: | 9
Comment | The content of | Paracetamol |  | in the sample as a percentage of the stated label claim is |  | 99.87%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Paracetamol Acetaminophen
Code: | P49-1
% age Purity: | 99.83
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.83
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 14567127 | 17211527.8209 | 15247245 | 16862313.9199
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 14561840 | 17205281.0608 | 15247148 | 16862206.6451
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 14579104 | 17225679.0306 | 15246471 | 16861457.9337
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 14569357 | 17214162.6375 | 15246954.6667 | 16861992.8329
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 23.55 |  | 25.16
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 23.55 |  | 25.16
Standard Dilution Factor | 500 | Purity correction (mg): | 23.51 |  | 25.1172
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.04702 |  | 0.050234
|  | Desired Concetration (mg/mL): | 0.055556
|  | Desired Weight as free base (mg): | 27.7778
|  | Desired Weight as salt (mg): | 27.7778
|  | Average Normalised Peak Area: | 17038077.7352
|  | RSD: | 1.133%
|  | n: | 6
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 5 | 1 | 15205409 | 446.2184 | 89.24%
1st Dilution Sample Final Volume (mL): | 50 | 2 | 14397555 | 422.5111 | 84.5%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 14614239 | 428.8699 | 85.77%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 15103866 | 443.2386 | 88.65%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 13786004 | 404.5645 | 80.91%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 15427518 | 452.7365 | 90.55%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 86.6%
Sample Dilution Factor | 9000 |  |  | RSD: | 4.137%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Paracetamol-Acetaminophen
Code: | P49-1
% age Purity: | 99.83
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.83
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 15445402 | 17343404.0631 | 16586326 | 17042682.9664
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 15470244 | 17371298.7624 | 16585833 | 17042176.402
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 15471974 | 17373241.3527 | 16556516 | 17012052.7727
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 15462540 | 17362648.0594 | 16576225 | 17032304.047
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 12.39 |  | 13.54
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 12.39 |  | 13.54
Standard Dilution Factor | 250 | Purity correction: | 12.3689 |  | 13.517
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.049476 |  | 0.054068
|  | Desired Concetration (mg/mL): | 0.055556
|  | Desired Weight as free base (mg): | 13.8889
|  | Desired Weight as salt (mg): | 13.8889
|  | Average Normalised Peak Area: | 17197476.0532
|  | RSD: | 1.056%
|  | n: | 6
Medium Volume (mL): | 900 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 10 | 1 | 14785731 | 429.8808 | 85.98%
1st Dilution Sample Final Volume (mL): | 100 | 2 | 14994212 | 435.9422 | 87.19%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 14883855 | 432.7337 | 86.55%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 14993926 | 435.9339 | 87.19%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 14514798 | 422.0037 | 84.4%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 14921274 | 433.8216 | 86.76%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 86.34%
Sample Dilution Factor | 9000 |  |  | RSD: | 1.22%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 86.47%
RSD: | 2.916%
n: | 12
Comment | The amount of | Paracetamol |  | dissolved as a percentage of the stated label claim is |  | 86.47%
| Name |  |  | Date |  | Signature
Analysed by: |  | RUTTO/JOYFRIDA
Reviewed By:
5