## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | EPLONE 25
Reference Substance: | Eplerenone
% age Purity: | 99.86
Weight (mg): | 25.37
Standard Conc (mg/mL): | 0.2537
2016-03-24
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 20265624 | 5478.25 | 1.4 | 2.68
2 | 20208151 | 5461.88 | 1.34 | 2.68
3 | 20228262 | 5482.48 | 1.37 | 2.68
4 | 20217948 | 5446.2 | 1.33 | 2.68
5 | 20152093 | 5496.99 | 1.39 | 2.68
6 | 20304051 | 5463.47 | 1.32 | 2.69
Averages: | 20229354.8333 | 5471.545 | 1.3583 | 2.6817
RSD: | 0.2563%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2500
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Eplerenone
% age Purity: | 99.86
Weight (mg): | 25
Standard Conc (mg/mL): | 0.025
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 9614445 | 7444.3 | 1.4 | 4.9
2 | 9553615 | 7452.3 | 1.3 | 4.9
3 | 9621428 | 7419.3 | 1.4 | 4.9
4 | 9611760 | 7422.8 | 1.4 | 4.9
5 | 9623489 | 7430.5 | 1.3 | 4.9
6 | 9578100 | 7407.8 | 1.3 | 4.9
Averages: | 9600472.8333 | 7429.5 | 1.35 | 4.9
RSD: | 0.2942%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2500
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: |  | RUTTO/MUTUA |  | 2016-05-04
Reviewed By:

## SST (S2)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | EPLONE 25
Reference Substance: | Eplerenone
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
2016-03-24
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages:
RSD:
n:
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution Repeat
Reference Substance: | Eplerenone
% age Purity: | 99.86
Weight (mg): | 25.01
Standard Conc (mg/mL): | 0.02501
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 6731097 | 5610.22 | 1.27 | 3.08
2 | 6738208 | 5595.68 | 1.25 | 3.08
3 | 6751980 | 5602.72 | 1.27 | 3.08
4 | 6715254 | 5582.09 | 1.3 | 3.07
5 | 6677804 | 5595.24 | 1.28 | 3.07
6 | 6701571 | 5612.09 | 1.26 | 3.08
Averages: | 9600472.8333 | 7429.5 | 1.35 | 4.9
RSD: | 0.2802%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2500
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: |  | RUTTO/MUTUA |  | 2016-05-04
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | EPLONE 25
Laboratory Ref No: |  | NDQD201512596
Active Ingredient: |  | Eplerenone
Label Claim: |  | Each film coated tablet contains: Eplerenone 25mg
Date Analysis Started: |  | 2015-12-08 12:13:53
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 154.29 | -0.7922%
|  | 154.94 | -0.3742%
|  | 155.68 | 0.1016%
|  | 153.87 | -1.062%
|  | 157.53 | 1.291%
|  | 155.85 | 0.2109%
|  | 156.51 | 0.6353%
|  | 154.03 | -0.9593%
|  | 155.13 | -0.2521%
|  | 155.44 | -0.05273%
|  | 155.19 | -0.2135%
|  | 153.47 | -1.319%
|  | 155.74 | 0.1402%
|  | 154.05 | -0.9465%
|  | 159.12 | 2.313%
|  | 155.92 | 0.2559%
|  | 158.13 | 1.677%
|  | 154.53 | -0.6379%
|  | 155.55 | 0.018%
|  | 155.47 | -0.03344%
| Total | 3110.44
| Average | 155.522
| Average | % Deviation from mean
| 155.522 | -7.5% | 143.8579
|  | 7.5% | 167.1862
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Eplerenone
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | EPLONE 25
Laboratory Ref No: | NDQD201512596 | 11
Active Ingredient: | Eplerenone
Label Claim: | Each film coated tablet contains: Eplerenone 25mg
Date Analysis Started: | 2016-03-24
Date Analysis Completed: | 2016-03-24
Analysis Data
Reference Substance: | Eplerenone
Code: | E10-2
% age Purity: | 99.86
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.86
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 25 | 1 | 20279645 | 20011900.1841 | 19420553 | 19731919.5331
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 20148750 | 19882733.3434 | 19408798 | 19719976.0671
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 20186840 | 19920320.4549 | 19299718 | 19609147.2054
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 20205078.3333 | 19938317.9942 | 19376356.3333 | 19687014.2686
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 25.37 |  | 24.64
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 25.37 |  | 24.64
Standard Dilution Factor | 100 | Purity correction (mg): | 25.3345 |  | 24.6055
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.25334 |  | 0.24606
|  | Desired Concetration (mg/mL): | 0.25
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Response: | 19812666.1314
|  | RSD: | 0.7578%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each film coated tablet contains: Eplerenone 25mg
Each Tablet contains | 25 | Eplerenone
Average Tablet Content Weight (mg): | 155.522
Initial Sample dilution (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 1 | Assay Smp A | 157.72 | 1 | 18686952 | 23.2509 | 93%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 18564575 | 23.0987 | 92.39%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 18670232 | 23.2301 | 92.92%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 155.88 | 1 | 18101394 | 22.7882 | 91.15%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 18062314 | 22.739 | 90.96%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 18106451 | 22.7946 | 91.18%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 100 | Assay Smp C | 156.05 | 1 | 18312508 | 23.0289 | 92.12%
Desired Sample Weight (mg) | 155.522 |  |  | 2 | 18248883 | 22.9489 | 91.8%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 18267412 | 22.9722 | 91.89%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 91.93%
[6 empty] | RSD: | 0.8178%
[6 empty] | n: | 9
Comment | The content of | Eplerenone |  | in the sample as a percentage of the stated label claim is |  | 91.93%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Eplerenone
Code: | E10-2
% age Purity: | 99.86
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.86
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 9505413 | 9752806.5932 | 9711468 | 9556881.9933
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 9546358 | 9794817.2524 | 9759988 | 9604629.6576
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 9551447 | 9800038.7017 | 9759819 | 9604463.3477
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 9534406 | 9782554.1824 | 9743758.3333 | 9588658.3329
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 24.4 |  | 25.44
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 24.4 |  | 25.44
Standard Dilution Factor | 1000 | Purity correction (mg): | 24.3658 |  | 25.4044
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.024366 |  | 0.025404
|  | Desired Concetration (mg/mL): | 0.025
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Peak Area: | 9685606.2576
|  | RSD: | 1.124%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 8262195 | 21.326 | 85.3%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 7070193 | 18.2492 | 73%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 8186828 | 21.1314 | 84.53%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 7903429 | 20.3999 | 81.6%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 9174972 | 23.682 | 94.73%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 7623263 | 19.6768 | 78.71%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 82.98%
Sample Dilution Factor | 1000 |  |  | RSD: | 8.784%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Eplerenone
Code: | E10-2
% age Purity: | 99.86
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.86
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 6719045 | 6725774.541 | 6552443 | 6684626.4069
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 6688912 | 6695611.3609 | 6522449 | 6654027.3335
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 6730615 | 6737356.129 | 6557852 | 6690144.5234
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 6712857.3333 | 6719580.677 | 6544248 | 6676266.0879
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 25.01 |  | 24.54
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 25.01 |  | 24.54
Standard Dilution Factor | 1000 | Purity correction: | 24.975 |  | 24.5056
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.024975 |  | 0.024506
|  | Desired Concetration (mg/mL): | 0.025
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Peak Area: | 6697923.3825
|  | RSD: | 0.4479%
|  | n: | 6
Medium Volume (mL): | 1000 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 4605524 | 17.1901 | 68.76%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 4928829 | 18.3969 | 73.59%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 5987414 | 22.348 | 89.39%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4961734 | 18.5197 | 74.08%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 4658045 | 17.3862 | 69.54%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 5509698 | 20.5649 | 82.26%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 76.27%
Sample Dilution Factor | 1000 |  |  | RSD: | 10.52%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 79.62%
RSD: | 10.18%
n: | 12
Comment | The amount of | Eplerenone |  | dissolved as a percentage of the stated label claim is |  | 79.62%
| Name |  |  | Date |  | Signature
Analysed by: | RUTTO/MUTUA
Reviewed By:
5