## SST
HPLC System Suitability Report
Analysis Data | Assay | Cefpodoxime R
Sample(s) | Aivpod-100 tablets
Reference Substance: | Cefpodoxime Proxetil
% age Purity: | 93.77
Weight (mg): | 24.17
Standard Conc (mg/mL): | Each uncoated dispersible tablet contains: Cefpodoxime Proxetil USP eq. to Cefpodoxime 100mg
2015-12-09 08:16:45
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3387326 | 13368.8 | 1.1 | 10.3
2 | 3407486 | 13413.1 | 1.1 | 10.3
3 | 3390908 | 13345.1 | 1.2 | 10.3
4 | 3408089 | 13376.3 | 1.1 | 10.3
5 | 3421453 | 13368.8 | 1.1 | 10.3
6 | 3434026 | 13364.5 | 1.1 | 10.3
Averages: | 3408214.6667 | 13372.7667 | 1.1167 | 10.3
RSD: | 0.5212%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Cefpodoxime S
Reference Substance: | Cefpodoxime Proxetil
% age Purity: | 93.77
Weight (mg): | 24.17
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3434638 | 13272.9 | 1.1 | 11.3
2 | 3455879 | 13283.5 | 1.1 | 11.3
3 | 3440807 | 13257.5 | 1.2 | 11.3
4 | 3461004 | 13292.5 | 1.1 | 11.3
5 | 3469199 | 13306.4 | 1.1 | 11.3
6 | 3481814 | 13298.7 | 1.2 | 11.3
Averages: | 3457223.5 | 13285.25 | 1.1333 | 11.3
RSD: | 0.5083%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | Aivpod-100 tablets
Laboratory Ref No: |  | NDQD201512616
Active Ingredient: |  | Cefpodoxime Proxetil
Label Claim: |  | Each uncoated dispersible tablet contains: Cefpodoxime Proxetil USP eq. to Cefpodoxime 100mg
Date Analysis Started: |  | 2015-12-09 08:16:45
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 295.69 | 1.396%
|  | 285.4 | -2.133%
|  | 295.93 | 1.478%
|  | 290.43 | -0.4077%
|  | 280.86 | -3.689%
|  | 293.34 | 0.5902%
|  | 289.88 | -0.5963%
|  | 283.99 | -2.616%
|  | 305.45 | 4.743%
|  | 289.79 | -0.6272%
|  | 291.26 | -0.1231%
|  | 291.8 | 0.06207%
|  | 308.97 | 5.95%
|  | 289.57 | -0.7026%
|  | 283.29 | -2.856%
|  | 287.51 | -1.409%
|  | 292.58 | 0.3295%
|  | 294.71 | 1.06%
|  | 288.44 | -1.09%
|  | 293.49 | 0.6416%
| Total | 5832.38
| Average | 291.619
| Average | % Deviation from mean
| 291.619 | -5% | 277.038
|  | 5% | 306.2
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Cefpodoxime
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Aivpod-100 tablets
Laboratory Ref No: | NDQD201512616 | 29
Active Ingredient: | Cefpodoxime Proxetil
Label Claim: | Each uncoated dispersible tablet contains: Cefpodoxime Proxetil USP eq. to Cefpodoxime 100mg
Date Analysis Started: | 2016-01-19
Date Analysis Completed: | 2016-01-29
Analysis Data
Reference Substance: | Cefpodoxime Proxetil
Code: | C68-1
% age Purity: | 93.77
% Water content: |  | If correction for water content is not needed please enter 0
% Purity corrected for water: | 93.77
Mwt of compound in free base form: | 427.458 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 557.5972 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.76661 | free base
Initial Standard dilution volume (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 6885792 | 7926298.9504 | 6200385 | 7906005.4539
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 6897717 | 7940025.9284 | 6259737 | 7981684.1796 |  | 0.5689%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 6932303 | 7979838.1933 | 6359380 | 8108737.2741
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 6905270.6667 | 7948721.024 | 6273167.3333 | 7998808.9692
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 24.17 |  | 21.82
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 18.5289 |  | 16.7274
Standard Dilution Factor | 625 | Purity correction (mg): | 17.3745 |  | 15.6853
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.027799 |  | 0.025096
|  | Desired Concentration (mg/mL): | 0.032
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 26.089
|  | Average Normalised Response: | 7973764.9966
|  | RSD: | 0.9096%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each uncoated dispersible tablet contains: Cefpodoxime Proxetil USP eq. to Cefpodoxime 100mg
Each Tablet contains | 100 | Cefpodoxime Proxetil
Average Tablet Content Weight (mg): | 291.619
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 4 | Assay Smp A | 184.44 | 1 | 6201846 | 98.3803 | 98.38%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 6170936 | 97.89 | 97.89%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 6194924 | 98.2705 | 98.27%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 183.57 | 1 | 6131531 | 97.7259 | 97.73%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 6143077 | 97.9099 | 97.91%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 6158799 | 98.1605 | 98.16%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 2500 | Assay Smp C | 184.44 | 1 | 6147415 | 97.5169 | 97.52%
Desired Sample Weight (mg): | 233.2952 |  |  | 2 | 6143761 | 97.4589 | 97.46%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 6134067 | 97.3051 | 97.31%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 97.8464 | 97.85%
[5 empty] | RSD: | 0.3851% | 0.3851%
[5 empty] | n: | 9 | 9
Comment: | The content of | Cefpodoxime Proxetil |  | in the sample as a percentage of the stated label claim is |  | 97.85%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Cefpodoxime Proxetil
Code: | C68-1
% age Purity: | 93.77
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 93.77
Mwt of compound in free base form: | 427.458 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 557.5972 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.76661 | free base
Initial Standard dilution volume (mL): | 1 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 1 | 1 |  | - |  | -
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 |  | - |  | - |  | #DIV/0!
2nd Dilution Standard Final Volume (mL): | 1 | 3 |  | - |  | -
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg):
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 0 |  | 0
Standard Dilution Factor | 1 | Purity correction: | 0 |  | 0
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0 |  | 0
|  | Desired Concentration (mg/mL): | 0.11111
|  | Desired Weight as free base (mg): | 0.11111
|  | Desired Weight as salt (mg): | 0.14494
|  | Average Normalised Peak Area: | #DIV/0!
|  | RSD: | #DIV/0!
|  | n: | 0
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 |  | - | -
1st Dilution Sample Final Volume (mL): | 1 | 2 |  | - | -
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 |  | - | -
2nd Dilution Sample Final Volume (mL): | 1 | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 |  | - | -
3rd Dilution Sample Final Volume (mL): | 1 | 6 |  | - | -
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | #DIV/0! | #DIV/0!
Sample Dilution Factor | 900 |  | RSD: | #DIV/0! | #DIV/0!
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 0 | 0
Comment: | The amount of | Cefpodoxime Proxetil |  | dissolved as a percentage of the stated label claim is |  | #DIV/0!
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5