## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | PROLERT CAPSULES
Reference Substance: | Fluoxetine HCl
% age Purity: | 99.52
Weight (mg): | 10.42
Standard Conc (mg/mL): | 0.1042
2016-11-05 18:07
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 23193406 | 7312.9 | 1.3 | 7
2 | 23299132 | 7351.4 | 1.3 | 7
3 | 23225878 | 7509.1 | 1.3 | 7.1
4 | 23300613 | 7389.9 | 1.3 | 7.1
5 | 23303326 | 7399.1 | 1.3 | 7.1
6 | 23324634 | 7437.5 | 1.3 | 7.1
Averages: | 23274498.1667 | 7399.9833 | 1.3 | 7.0667
RSD: | 0.2239%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | First Dissolution (S1)
Reference Substance: | Fluoxetine HCl
% age Purity: | 99.52
Weight (mg): | 19.3
Standard Conc (mg/mL): | 0.0193
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 17243788 | 8400.9 | 1.2 | 3.2
2 | 17251984 | 8411 | 1.2 | 3.2
3 | 17261435 | 8429.6 | 1.2 | 3.2
4 | 17241835 | 8427.5 | 1.2 | 3.2
5 | 17237212 | 8446.3 | 1.2 | 3.2
6 | 17259445 | 8459 | 1.2 | 3.2
Averages: | 17249283.1667 | 8429.05 | 1.2 | 3.2
RSD: | 0.05737%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST (S2)
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | PROLERT CAPSULES
Reference Substance: | Fluoxetine HCl
% age Purity: | 99.52
Weight (mg): | 10.42
Standard Conc (mg/mL): | 0.1042
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 23193406 | 7312.9 | 1.3 | 7
2 | 23299132 | 7351.4 | 1.3 | 7
3 | 23225878 | 7509.1 | 1.3 | 7.1
4 | 23300613 | 7389.9 | 1.3 | 7.1
5 | 23303326 | 7399.1 | 1.3 | 7.1
6 | 23324634 | 7437.5 | 1.3 | 7.1
Averages: | 23274498.1667 | 7399.9833 | 1.3 | 7.0667
RSD: | 0.2239%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 1100
The Assymetry of all peaks were below 2.0
Analysis Data | Repeat Dissolution (S2)
Reference Substance: | Fluoxetine HCl
% age Purity: | 99.52
Weight (mg): | 19.82
Standard Conc (mg/mL): | 0.01982
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 18341580 | 6619.35 | 1.22 | 2.32
2 | 18164323 | 6686.49 | 1.23 | 2.32
3 | 18116217 | 6739.07 | 1.2 | 2.33
4 | 18121545 | 6722.22 | 1.24 | 2.32
5 | 18213059 | 6759.25 | 1.2 | 2.33
6 | 17923168 | 6744.63 | 1.21 | 2.33
Averages: | 18146648.6667 | 6711.835 | 1.2167 | 2.325
RSD: | 0.7565%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 1100
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | PROLERT CAPSULES
Laboratory Ref No: |  | NDQD201601651
Active Ingredient: |  | Fluoxetine HCl
Label Claim: |  | Each capsule contains Fluoxetine 20 mg as Fluoxetine Hydrochloride BP
Date Analysis Started: |  | 2016-01-12 14:45:04
Date Analysis Completed: |  | 2016-04-22 07:12:38
Analysis Data |  | Uniformity of weight
Capsule No. | Intact Capsule (mg) | Empty Shell (mg) | Capsule Content (mg) | % Deviation
1 | 348.64 | 64.31 | 284.33 | -1.79%
2 | 353.82 | 63.54 | 290.28 | 0.2648%
3 | 365.22 | 64.08 | 301.14 | 4.016%
4 | 357.32 | 64.1 | 293.22 | 1.28%
5 | 359.96 | 66.15 | 293.81 | 1.484%
6 | 345.34 | 61.49 | 283.85 | -1.956%
7 | 354.16 | 63.05 | 291.11 | 0.5514%
8 | 343.76 | 61.68 | 282.08 | -2.568%
9 | 375.06 | 64.76 | 310.3 | 7.18%
10 | 342.04 | 64.15 | 277.89 | -4.015%
11 | 359.82 | 67.02 | 292.8 | 1.135%
12 | 349.08 | 63.73 | 285.35 | -1.438%
13 | 363.19 | 63.88 | 299.31 | 3.384%
14 | 342.71 | 62.35 | 280.36 | -3.162%
15 | 354.76 | 59.97 | 294.79 | 1.823%
16 | 353.55 | 61.95 | 291.6 | 0.7207%
17 | 347.26 | 63.18 | 284.08 | -1.877%
18 | 346.49 | 60.15 | 286.34 | -1.096%
19 | 334.75 | 59.91 | 274.84 | -5.068%
20 | 355.82 | 63.03 | 292.79 | 1.132%
Total | 7052.75 | 1262.48 | 5790.27
Average | 352.6375 | 63.124 | 289.5135
| Average | % Deviation from mean
| 289.5135 | -10% | 260.5622
|  | 10% | 318.4649
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Fluoxetine HCl
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | PROLERT CAPSULES
Laboratory Ref No: | NDQD201601651 | 11
Active Ingredient: | Fluoxetine HCl
Label Claim: | Each capsule contains Fluoxetine 20 mg as Fluoxetine Hydrochloride BP
Date Analysis Started: | 2016-04-13 11:16
Date Analysis Completed: | 2016-05-11 20:31
Analysis Data
Reference Substance: | Fluoxetine HCl
Code: | F18-1
% age Purity: | 99.52
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.52
Mwt of compound in free base form: | 309.3231 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 345.7871 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.89455 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 23309156 | 25127241.6112 | 22403701 | 25317415.5374
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 23160064 | 24966520.6179 | 22039918 | 24906320.7198
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 23218776 | 25029812.0819 | 22050390 | 24918154.6563
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 23229332 | 25041191.437 | 22164669.6667 | 25047296.9712
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 10.42 |  | 9.94
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 9.3212 |  | 8.8918
Standard Dilution Factor | 100 | Purity correction (mg): | 9.2764 |  | 8.8491
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.092764 |  | 0.088491
|  | Desired Concetration (mg/mL): | 0.1
|  | Desired Weight as free base (mg): | 10
|  | Desired Weight as salt (mg): | 11.1788
|  | Average Normalised Response: | 25044244.2041
|  | RSD: | 0.6255%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample |  |  |  |  | 97.55%
Label Claim: | Each capsule contains Fluoxetine 20 mg as Fluoxetine Hydrochloride BP
Each Capsule contains | 20 | Fluoxetine as Fluoxetine HCl
Average Capsule Content Weight (mg): | 289.5135
Initial Sample dilution (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 1 | Assay Smp A | 144.6 | 1 | 24606234 | 19.6715 | 98.36%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 24524956 | 19.6065 | 98.03%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 24509973 | 19.5946 | 97.97%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 144.02 | 1 | 24377000 | 19.5667 | 97.83%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 24653757 | 19.7889 | 98.94%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 24559828 | 19.7135 | 98.57%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 100 | Assay Smp C | 144.21 | 1 | 24156676 | 19.3643 | 96.82%
Desired Sample Weight (mg) | 144.7568 |  |  | 2 | 24224595 | 19.4188 | 97.09%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 24338142 | 19.5098 | 97.55%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 97.91%
[6 empty] | RSD: | 0.6954%
[6 empty] | n: | 9
Comment | The content of | Fluoxetine |  | in the sample as a percentage of the stated label claim is |  | 97.91%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Fluoxetine HCl
Code: | F18-1
% age Purity: | 99.52
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.52
Mwt of compound in free base form: | 309.3231 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 345.7871 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.89455 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 17242406 | 22300470.7319 | 18372459 | 21974464.9967
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 17231629 | 22286532.2958 | 19325880 | 23114808.616
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 17188736 | 22231056.622 | 18319382 | 21910982.0041
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 17220923.6667 | 22272686.5499 | 18672573.6667 | 22333418.5389
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 19.3 |  | 20.87
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 17.2648 |  | 18.6692
Standard Dilution Factor | 1000 | Purity correction (mg): | 17.1819 |  | 18.5796
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.017182 |  | 0.01858
|  | Desired Concetration (mg/mL): | 0.022222
|  | Desired Weight as free base (mg): | 22.2222
|  | Desired Weight as salt (mg): | 24.8418
|  | Average Normalised Peak Area: | 22303052.5444
|  | RSD: | 1.93%
|  | n: | 6 |  | 99.06%
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 20455849 | 18.3435 | 91.72%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 18656677 | 16.7302 | 83.65%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 22352136 | 20.044 | 100.2%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 19035728 | 17.0701 | 85.35%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 15366642 | 13.7799 | 68.9%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 22093434 | 19.812 | 99.06%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 88.15%
Sample Dilution Factor | 900 |  |  | RSD: | 13.19%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Fluoxetine HCl
Code: | F18-1
% age Purity: | 99.52
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.52
Mwt of compound in free base form: | 309.3231 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 345.7871 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.89455 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 18184373 | 22901722.7851 | 19306631 | 22992632.0998
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 18092392 | 22785880.2777 | 19188757 | 22852253.7233
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 18072648 | 22761014.3329 | 19247636 | 22922373.8383
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 18116471 | 22816205.7986 | 19247674.6667 | 22922419.8871
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.82 |  | 20.96
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 17.7299 |  | 18.7497
Standard Dilution Factor | 1000 | Purity correction: | 17.6448 |  | 18.6597
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.017645 |  | 0.01866
|  | Desired Concetration (mg/mL): | 0.022222
|  | Desired Weight as free base (mg): | 22.2222
|  | Desired Weight as salt (mg): | 24.8418
|  | Average Normalised Peak Area: | 22869312.8428
|  | RSD: | 0.3815%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 20083352 | 17.5636 | 87.82%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 20418063 | 17.8563 | 89.28%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 21251777 | 18.5854 | 92.93%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 19258515 | 16.8422 | 84.21%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 20389601 | 17.8314 | 89.16%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 20785502 | 18.1776 | 90.89%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 89.05%
Sample Dilution Factor | 900 |  |  | RSD: | 3.309%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 88.6%
RSD: | 9.144%
n: | 12
Comment | The amount of | Fluoxetine HCl |  | dissolved as a percentage of the stated label claim is |  | 88.6%
| Name |  |  | Date |  | Signature
Analysed by: | KIPKORIR / MICHAEL
Reviewed By:
5