## SST-S1
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | STODIL CAPSULES
Reference Substance: | Loperamide Hydrochloride
% age Purity: | 99.89
Weight (mg): | 20.04
Standard Conc (mg/mL): | 0.01002
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4471809 | 8288.91 | 1.35 | 6.41
2 | 4482047 | 8286.71 | 1.32 | 6.4
3 | 4480540 | 8281.59 | 1.36 | 6.39
4 | 4483172 | 8308.02 | 1.2 | 6.39
5 | 4481380 | 8332.74 | 1.32 | 6.38
6 | 4482038 | 8357.01 | 1.33 | 6.37
Averages: | 4480164.3333 | 8309.1633 | 1.3133 | 6.39
RSD: | 0.09339%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 1900
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Loperamide Hydrochloride
% age Purity: | 99.89
Weight (mg): | 20.4
Standard Conc (mg/mL): | 0.00408
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4786863 | 8263.72 | 1.36 | 5.51
2 | 4775097 | 8269.03 | 1.37 | 5.51
3 | 4770340 | 8287.84 | 1.35 | 5.51
4 | 4775234 | 8260.84 | 1.38 | 5.51
5 | 4760178 | 8323.14 | 1.35 | 5.51
6 | 4734218 | 8149.88 | 1.38 | 5.5
Averages: | 4766988.3333 | 8259.075 | 1.365 | 5.5083
RSD: | 0.3823%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 1900
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO/JOYFRIDA |  |  | 2016-01-08
Reviewed By:

## SST-S2
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | STODIL CAPSULES
Reference Substance: | Loperamide Hydrochloride
% age Purity: | 99.89
Weight (mg): | 20.04
Standard Conc (mg/mL): | 0.01002
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4471809 | 8288.91 | 1.35 | 6.41
2 | 4482047 | 8286.71 | 1.32 | 6.4
3 | 4480540 | 8281.59 | 1.36 | 6.39
4 | 4483172 | 8308.02 | 1.2 | 6.39
5 | 4481380 | 8332.74 | 1.32 | 6.38
6 | 4482038 | 8357.01 | 1.33 | 6.37
Averages: | 4480164.3333 | 8309.1633 | 1.3133 | 6.39
RSD: | 0.09339%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 1900
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Loperamide Hydrochloride
% age Purity: | 99.89
Weight (mg): | 19.15
Standard Conc (mg/mL): | 0.00383
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 4449186 | 8524.67 | 1.33 | 5.44
2 | 4482505 | 8555.8 | 1.31 | 5.44
3 | 4589953 | 8564.06 | 1.32 | 5.44
4 | 4629632 | 8620.2 | 1.33 | 5.44
5 | 4604368 | 8635.37 | 1.31 | 5.44
6 | 4595421 | 8627.72 | 1.34 | 5.44
Averages: | 4558510.8333 | 8587.97 | 1.3233 | 5.44
RSD: | 1.619%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is NLT 1900
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | RUTTO/JOYFRIDA |  |  | 2016-01-08
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | STODIL CAPSULES
Laboratory Ref No: |  | NDQD201605936
Active Ingredient: |  | Loperamide
Label Claim: |  | Each capsule contains Laperamide hydrochloride USP 2 mg
Date Analysis Started: |  | 2016-05-13 08:55:55
Date Analysis Completed: |  | 2016-07-27 09:12:22
Analysis Data |  | Uniformity of weight
Capsule No. | Intact Capsule (mg) | Empty Shell (mg) | Capsule Content (mg) | % Deviation
1 | 204.83 | 38.17 | 166.66 | 8.904%
2 | 201.11 | 39.77 | 161.34 | 5.427%
3 | 195.28 | 36.57 | 158.71 | 3.709%
4 | 177.02 | 38.61 | 138.41 | -9.556%
5 | 199.76 | 37.34 | 162.42 | 6.133%
6 | 200.7 | 37.53 | 163.17 | 6.623%
7 | 176.2 | 41.93 | 134.27 | -12.26%
8 | 199.17 | 38.4 | 160.77 | 5.055%
9 | 193.62 | 39.35 | 154.27 | 0.8073%
10 | 185.61 | 37.06 | 148.55 | -2.93%
11 | 175.62 | 39.44 | 136.18 | -11.01%
12 | 188.59 | 40.78 | 147.81 | -3.414%
13 | 192.88 | 40.03 | 152.85 | -0.1206%
14 | 194.09 | 38.4 | 155.69 | 1.735%
15 | 193.56 | 40.63 | 152.93 | -0.06829%
16 | 198.11 | 41.45 | 156.66 | 2.369%
17 | 186.93 | 39.36 | 147.57 | -3.571%
18 | 186.31 | 40.7 | 145.61 | -4.852%
19 | 193.36 | 39.07 | 154.29 | 0.8204%
20 | 199.62 | 37.09 | 162.53 | 6.205%
Total | 3842.37 | 781.68 | 3060.69
Average | 192.1185 | 39.084 | 153.0345
| Average | % Deviation from mean
| 153.0345 | -10% | 137.7311
|  | 10% | 168.338
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## Loperamide Hydrochloride
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | STODIL CAPSULES
Laboratory Ref No: | NDQD201605936 | 29
Active Ingredient: | Loperamide
Label Claim: | Each capsule contains Laperamide hydrochloride USP 2 mg
Date Analysis Started: | 2016-07-26
Date Analysis Completed: | 2016-07-27
Analysis Data
Reference Substance: | Loperamide Hydrochloride
Code: | L38-1
% age Purity: | 99.89
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.89
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4469901 | 4465891.5226 | 5047597 | 4413236.2192
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4475384 | 4471369.6044 | 5041030 | 4407494.5322 |  | 1.497%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4462045 | 4458042.5694 | 5031372 | 4399050.3091
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4469110 | 4465101.2321 | 5039999.6667 | 4406593.6868
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 20.04 |  | 22.9
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 20.04 |  | 22.9
Standard Dilution Factor | 2000 | Purity correction (mg): | 20.018 |  | 22.8748
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.010009 |  | 0.011437
|  | Desired Concentration (mg/mL): | 0.01
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 4435847.4595
|  | RSD: | 0.7358%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each capsule contains Laperamide hydrochloride USP 2 mg
Each Tablet/Capsule contains | 2 | Loperamide
Average Tablet/Capsule Content Weight (mg): | 153.0345
Initial Sample dilution Volume (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 762.33 | 1 | 4191830 | 1.897 | 94.85%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 4226347 | 1.9126 | 95.63%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 4213676 | 1.9069 | 95.35%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 756.34 | 1 | 4198681 | 1.9152 | 95.76%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 4192204 | 1.9122 | 95.61%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 4196737 | 1.9143 | 95.71%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 767.31 | 1 | 4333302 | 1.9483 | 97.42%
Desired Sample Weight (mg): | 765.1725 |  |  | 2 | 4328411 | 1.9461 | 97.31%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 4328593 | 1.9462 | 97.31%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 1.9221 | 96.11%
[5 empty] | RSD: | 1.008% | 1.008%
[5 empty] | n: | 9 | 9
Comment: | The content of | Loperamide |  | in the sample as a percentage of the stated label claim is |  | 96.11%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Loperamide Hydrochloride
Code: | L38-1
% age Purity: | 99.89
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.89
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4827411 | 4737967.6468 | 4453977 | 4683699.3382
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4823164 | 4733799.3361 | 4443482 | 4672663.0386 |  | 1.101%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4814914 | 4725702.194 | 4444392 | 4673619.9736
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4821829.6667 | 4732489.7256 | 4447283.6667 | 4676660.7834
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 20.4 |  | 19.04
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 20.4 |  | 19.04
Standard Dilution Factor | 5000 | Purity correction: | 20.3776 |  | 19.0191
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.0040755 |  | 0.0038038
|  | Desired Concentration (mg/mL): | 0.004
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Peak Area: | 4704575.2545
|  | RSD: | 0.6605%
|  | n: | 6
Medium Volume (mL): | 500 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 4201954 | 1.7863 | 89.32%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 4190512 | 1.7815 | 89.07%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 3995290 | 1.6985 | 84.92%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4294153 | 1.8255 | 91.28%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 3839577 | 1.6323 | 81.61%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 3826151 | 1.6266 | 81.33%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 1.7251 | 86.26%
Sample Dilution Factor | 500 |  | RSD: | 4.922% | 4.922%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Loperamide |  | dissolved as a percentage of the stated label claim is |  | 86.26%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Loperamide Hydrochloride-S2
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | STODIL CAPSULES
Laboratory Ref No: | NDQD201605936 | 11
Active Ingredient: | Loperamide
Label Claim: | Each capsule contains Laperamide hydrochloride USP 2 mg
Date Analysis Started: | 2016-07-26
Date Analysis Completed: | 2016-07-27
Analysis Data
Reference Substance: | Loperamide Hydrochloride
Code: | L38-1
% age Purity: | 99.89
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.89
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4469901 | 4465891.5226 | 5047597 | 4413236.2192
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4475384 | 4471369.6044 | 5041030 | 4407494.5322
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4462045 | 4458042.5694 | 5031372 | 4399050.3091
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4469110 | 4465101.2321 | 5039999.6667 | 4406593.6868
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 20.04 |  | 22.9
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 20.04 |  | 22.9
Standard Dilution Factor | 2000 | Purity correction (mg): | 20.018 |  | 22.8748
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.010009 |  | 0.011437
|  | Desired Concetration (mg/mL): | 0.01
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Response: | 4435847.4595
|  | RSD: | 0.7358%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each capsule contains Laperamide hydrochloride USP 2 mg
Each Tablet contains | 2 | Loperamide
Average Tablet Content Weight (mg): | 153.0345
Initial Sample dilution (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 5 | Assay Smp A | 762.33 | 1 | 4191830 | 1.897 | 94.85%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 4226347 | 1.9126 | 95.63%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 4213676 | 1.9069 | 95.35%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 756.34 | 1 | 4198681 | 1.9152 | 95.76%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 4192204 | 1.9122 | 95.61%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 4196737 | 1.9143 | 95.71%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 767.31 | 1 | 4333302 | 1.9483 | 97.42%
Desired Sample Weight (mg) | 765.1725 |  |  | 2 | 4328411 | 1.9461 | 97.31%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 4328593 | 1.9462 | 97.31%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 96.11%
[6 empty] | RSD: | 1.008%
[6 empty] | n: | 9
Comment | The content of | Loperamide |  | in the sample as a percentage of the stated label claim is |  | 96.11%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Loperamide Hydrochloride
Code: | L38-1
% age Purity: | 99.89
% Water content: | 0 | If correction for water content is NOT needed please enter 0
% Purity corrected for water: | 99.89
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4827411 | 4737967.6468 | 4453977 | 4683699.3382
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4823164 | 4733799.3361 | 4443482 | 4672663.0386
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4814914 | 4725702.194 | 4444392 | 4673619.9736
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4821829.6667 | 4732489.7256 | 4447283.6667 | 4676660.7834
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 20.4 |  | 19.04
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 20.4 |  | 19.04
Standard Dilution Factor | 5000 | Purity correction (mg): | 20.3776 |  | 19.0191
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.0040755 |  | 0.0038038
|  | Desired Concetration (mg/mL): | 0.004
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Peak Area: | 4704575.2545
|  | RSD: | 0.6605%
|  | n: | 6
Medium Volume (mL): | 500 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 4201954 | 1.7863 | 89.32%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 4190512 | 1.7815 | 89.07%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 3995290 | 1.6985 | 84.92%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4294153 | 1.8255 | 91.28%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 3839577 | 1.6323 | 81.61%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 3826151 | 1.6266 | 81.33%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 86.26%
Sample Dilution Factor | 500 |  |  | RSD: | 4.922%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Repeat Determination of Active Ingredient Dissolved
Reference Substance: | Loperamide Hydrochloride
Code: | L38-1
% age Purity: | 99.89
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.89
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 4265111 | 4459329.2831 | 4902600 | 4589059.1853
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 4260682 | 4454698.6019 | 4893227 | 4580285.626
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 4251161 | 4444744.0487 | 4887944 | 4575340.4949
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 4258984.6667 | 4452923.9779 | 4894590.3333 | 4581561.7687
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 19.15 |  | 21.39
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 19.15 |  | 21.39
Standard Dilution Factor | 5000 | Purity correction: | 19.1289 |  | 21.3665
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.0038258 |  | 0.0042733
|  | Desired Concetration (mg/mL): | 0.004
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Peak Area: | 4517242.8733
|  | RSD: | 1.566%
|  | n: | 6
Medium Volume (mL): | 500 | tablet No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 3589380 | 1.5892 | 79.46%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 4081088 | 1.8069 | 90.34%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 3955972 | 1.7515 | 87.57%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 4062724 | 1.7988 | 89.94%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 4051253 | 1.7937 | 89.68%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 3528039 | 1.562 | 78.1%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  |  | Average: | 85.85%
Sample Dilution Factor | 500 |  |  | RSD: | 6.495%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | n: | 6
Analysis Data: | Dissolution Result Summary
Average: | 86.05%
RSD: | 5.496%
n: | 12
Comment | The amount of | Loperamide |  | dissolved as a percentage of the stated label claim is |  | 86.05%
| Name |  |  | Date |  | Signature
Analysed by: |  | RUTTO/JOYFRIDA
Reviewed By:
5