## SST 3TC
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | NDQD2016061020
Reference Substance: | Lamivudine
% age Purity: | 100.4
Weight (mg): | 20.59
Standard Conc (mg/mL): | 0.1
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 57316657 | 6851.43 | 1.28 | 2.98
2 | 57592703 | 6852.52 | 1.3 | 2.98
3 | 57461813 | 6836.82 | 1.3 | 2.98
4 | 57444491 | 6840.24 | 1.28 | 2.98
5 | 57576504 | 6781.34 | 1.27 | 2.98
6 | 57172787 | 6794 | 1.31 | 2.98
Averages: | 57427492.5 | 6826.0583 | 1.29 | 2.98
RSD: | 0.2787%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Lamivudine
% age Purity: | 100.4
Weight (mg): | 20.59
Standard Conc (mg/mL): | 0.33
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 189121539 | 3668.96 | 1.4 | 2.99
2 | 189072227 | 3670.83 | 1.4 | 2.99
3 | 189292506 | 3665.92 | 1.43 | 2.99
4 | 189331486 | 3674.35 | 1.38 | 2.99
5 | 189639451 | 3656.48 | 1.42 | 2.99
6 | 189157722 | 3686.98 | 1.4 | 2.99
Averages: | 189269155.1667 | 3670.5867 | 1.405 | 2.99
RSD: | 0.1094%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## SST TDF
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | NDQD2016061020
Reference Substance: | Tenofovir Disoproxil Fumarate
% age Purity: | 98.8
Weight (mg): | 21.22
Standard Conc (mg/mL): | 0.1
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min) | Peak Resolution (USP)
1 | 57033835 | 9850.69 | 1.21 | 5.26 | 12.69
2 | 57280150 | 9583.6 | 1.22 | 5.25 | 12.67
3 | 57183938 | 9554.54 | 1.2 | 5.25 | 12.65
4 | 57141384 | 9557.98 | 1.19 | 5.25 | 12.65
5 | 57262716 | 9507.27 | 1.19 | 5.25 | 12.61
6 | 57172787 | 9541.48 | 1.21 | 5.25 | 12.63
Averages: | 57179135 | 9599.26 | 1.2033 | 5.2517 | 12.65
RSD: | 0.1743%
n: | 5
Summary: | The RSD of the peak areas for five replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | T enofovir Disoproxil Fumarate
% age Purity: | 98.8
Weight (mg): | 21.22
Standard Conc (mg/mL): | 0.333
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min) | Peak Resolution (USP)
1 | 187662528 | 6905.43 | 1.32 | 5.24 | 9.99
2 | 187616077 | 6906.15 | 1.33 | 5.24 | 9.99
3 | 187590542 | 6932.44 | 1.3 | 5.25 | 10.01
4 | 187637618 | 6900.86 | 1.32 | 5.24 | 9.99
5 | 187610848 | 6944.85 | 1.29 | 5.25 | 10.01
6 | 187437895 | 6929.26 | 1.32 | 5.24 | 10.01
Averages: | 187592584.6667 | 6919.8317 | 1.3133 | 5.2433 | 10
RSD: | 0.04247%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | LAMIVUDINE 300MG AND TENOFOVIR DISOPROXIL FUMARATE 300MG TABLETS
Laboratory Ref No: |  | NDQD2016061020
Active Ingredient: |  | Lamivudine and Tenofovir Disoproxil Fumarate
Label Claim: |  | Each tablet contains Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300mg
Date Analysis Started: |  | 2016-06-10 12:20:52
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1097.33 | -1.28%
|  | 1092.17 | -1.744%
|  | 1071.64 | -3.591%
|  | 1108.57 | -0.2691%
|  | 1084.58 | -2.427%
|  | 1094.51 | -1.534%
|  | 1114.41 | 0.2563%
|  | 1114.05 | 0.2239%
|  | 1121.86 | 0.9265%
|  | 1128.87 | 1.557%
|  | 1125.63 | 1.266%
|  | 1131.28 | 1.774%
|  | 1114.44 | 0.259%
|  | 1120.46 | 0.8006%
|  | 1094.7 | -1.517%
|  | 1113.76 | 0.1978%
|  | 1113.26 | 0.1528%
|  | 1136 | 2.199%
|  | 1120.33 | 0.7889%
|  | 1133.37 | 1.962%
| Total | 22231.22
| Average | 1111.561
| Average | % Deviation from mean
| 1111.561 | -5% | 1055.9829
|  | 5% | 1167.139
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## 3TC
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | LAMIVUDINE AND TENOFOVIR DISOPROXIL FUMARATE TABLETS
Laboratory Ref No: | NDQD2016061020 | 29
Active Ingredient: | Lamivudine
Label Claim: | Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300 mg
Date Analysis Started: | 2016-07-01
Date Analysis Completed: | 2016-07-05
Analysis Data
Reference Substance: | Lamivudine
Code: | L3-7
% age Purity: | 100.4
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 100.4
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 57767576 | 69860886.7106 | 56494703 | 72325995.5038
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 57892958 | 70012516.7131 | 56573689 | 72427115.4457 |  | 3.297%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 57896563 | 70016876.3992 | 56601404 | 72462596.9131
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 57852365.6667 | 69963426.6076 | 56556598.6667 | 72405235.9542
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 20.59 |  | 19.45
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 20.59 |  | 19.45
Standard Dilution Factor | 250 | Purity correction (mg): | 20.6724 |  | 19.5278
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.082689 |  | 0.078111
|  | Desired Concentration (mg/mL): | 0.1
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Response: | 71184331.2809
|  | RSD: | 1.882%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300 mg
Each Tablet contains | 300 | Lamivudine
Average Tablet Content Weight (mg): | 1111.561
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 183.43 | 1 | 64538921 | 274.7073 | 91.57%
1st Dilution Sample Final Volume (mL): | 25 |  |  | 2 | 64682510 | 275.3185 | 91.77%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 64570055 | 274.8398 | 91.61%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 185.29 | 1 | 66328253 | 279.4895 | 93.16%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 66453648 | 280.0179 | 93.34%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 66369529 | 279.6634 | 93.22%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 185.69 | 1 | 66471530 | 279.4899 | 93.16%
Desired Sample Weight (mg): | 185.2602 |  |  | 2 | 66229640 | 278.4728 | 92.82%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 66493467 | 279.5821 | 93.19%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 277.9535 | 92.65%
[5 empty] | RSD: | 0.8244% | 0.8244%
[5 empty] | n: | 9 | 9
Comment: | The content of | Lamivudine |  | in the sample as a percentage of the stated label claim is |  | 92.65%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Lamivudine
Code: | L3-7
% age Purity: | 100.4
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 100.4
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 15 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 189349697 | 254432188.9497 | 184865269 | 262965943.9932
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 189138009 | 254147740.4289 | 185212020 | 263459188.0977 |  | 3.319%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 189043478 | 254020717.626 | 184867814 | 262969564.1883
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 189177061.3333 | 254200215.6682 | 184981701 | 263131565.4264
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 20.59 |  | 19.45
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 20.59 |  | 19.45
Standard Dilution Factor | 83.3333 | Purity correction: | 20.6724 |  | 19.5278
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.24807 |  | 0.23433
|  | Desired Concentration (mg/mL): | 0.33333
|  | Desired Weight as free base (mg): | 27.7778
|  | Desired Weight as salt (mg): | 27.7778
|  | Average Normalised Peak Area: | 258665890.5473
|  | RSD: | 1.893%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 231609196 | 268.6197 | 89.54%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 233124106 | 270.3767 | 90.13%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 225400195 | 261.4185 | 87.14%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 228159316 | 264.6186 | 88.21%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 232026337 | 269.1035 | 89.7%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 227076054 | 263.3622 | 87.79%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 266.2499 | 88.75%
Sample Dilution Factor | 900 |  | RSD: | 1.356% | 1.356%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Lamivudine |  | dissolved as a percentage of the stated label claim is |  | 88.75%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## TDF
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | LAMIVUDINE AND TENOFOVIR DISOPROXIL FUMARATE TABLETS
Laboratory Ref No: | NDQD2016061020 | 29
Active Ingredient: | Tenofovir Disoproxil Fumarate
Label Claim: | Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300 mg
Date Analysis Started: | 2016-06-06
Date Analysis Completed: | 2016-06-07
Analysis Data
Reference Substance: | Tenofovir Disoproxil Fumarate
Code: | T11-8
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 57347458 | 56986319.7357 | 55969832 | 57263492.1664
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 57434558 | 57072871.2346 | 56039660 | 57334934.1377 |  | 0.4773%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 57460594 | 57098743.2762 | 56103892 | 57400650.7657
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 57414203.3333 | 57052644.7488 | 56037794.6667 | 57333025.6899
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 21.22 |  | 20.61
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 21.22 |  | 20.61
Standard Dilution Factor | 208.3333 | Purity correction (mg): | 20.9654 |  | 20.3627
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.10063 |  | 0.097741
|  | Desired Concentration (mg/mL): | 0.1
|  | Desired Weight as free base (mg): | 20.8333
|  | Desired Weight as salt (mg): | 20.8333
|  | Average Normalised Response: | 57192835.2194
|  | RSD: | 0.2865%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Lamivudine 300mg and Tenofovir Disoproxil Fumarate 300 mg
Each Tablet contains | 300 | Tenofovir Disoproxil Fumarate
Average Tablet Content Weight (mg): | 1111.561
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 183.43 | 1 | 54525632 | 288.8631 | 96.29%
1st Dilution Sample Final Volume (mL): | 25 |  |  | 2 | 54648137 | 289.5121 | 96.5%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 54501721 | 288.7364 | 96.25%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 185.29 | 1 | 55384821 | 290.4694 | 96.82%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 55412566 | 290.615 | 96.87%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 55363686 | 290.3586 | 96.79%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 185.69 | 1 | 56091071 | 293.5397 | 97.85%
Desired Sample Weight (mg): | 185.2602 |  |  | 2 | 56018921 | 293.1622 | 97.72%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 56055205 | 293.352 | 97.78%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 290.9565 | 96.99%
[5 empty] | RSD: | 0.6587% | 0.6587%
[5 empty] | n: | 9 | 9
Comment: | The content of | Tenofovir Disoproxil Fumarate |  | in the sample as a percentage of the stated label claim is |  | 96.99%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Tenofovir Disoproxil Fumarate
Code: | T11-8
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 25 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 187567235 | 186386054.3853 | 182816947 | 187042491.2552
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 187437200 | 186256838.264 | 182975573 | 187204783.6614 |  | 0.3976%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 187465696 | 186285154.8141 | 182744850 | 186968727.8394
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 187490043.6667 | 186309349.1545 | 182845790 | 187072000.9187
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 21.22 |  | 20.61
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 21.22 |  | 20.61
Standard Dilution Factor | 62.5 | Purity correction: | 20.9654 |  | 20.3627
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.33545 |  | 0.3258
|  | Desired Concentration (mg/mL): | 0.33333
|  | Desired Weight as free base (mg): | 20.8333
|  | Desired Weight as salt (mg): | 20.8333
|  | Average Normalised Peak Area: | 186690675.0366
|  | RSD: | 0.2286%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 176384659 | 283.4389 | 94.48%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 184649427 | 296.7198 | 98.91%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 174598411 | 280.5685 | 93.52%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 174053876 | 279.6935 | 93.23%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 179185927 | 287.9403 | 95.98%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 177365983 | 285.0158 | 95.01%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 285.5628 | 95.19%
Sample Dilution Factor | 900 |  | RSD: | 2.183% | 2.183%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Tenofovir Disoproxil Fumarate |  | dissolved as a percentage of the stated label claim is |  | 95.19%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5