## Uniformity (2)
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | TENOFOVIR DISOPROXIL FUMARATE/ LAMIVUDINE/ EFAVIRENZ TABLETS 300 MG/300 MG /600 MG
Laboratory Ref No: |  | NDQD2016061054
Active Ingredient: |  | Tenofovir Disoproxil Fumarate , Lamivudine & Efavirenz
Label Claim: |  | Tenofovir Disoproxil Fumarate 300mg, Lamivudine 300mg, Efavirenz 600mg
Date Analysis Started: |  | 2016-06-10 15:30:25
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1770.81 | 0.2177%
|  | 1790.14 | 1.312%
|  | 1773.48 | 0.3689%
|  | 1771.75 | 0.2709%
|  | 1763.79 | -0.1795%
|  | 1778.39 | 0.6467%
|  | 1755.95 | -0.6232%
|  | 1771.83 | 0.2755%
|  | 1769.88 | 0.1651%
|  | 1779.39 | 0.7033%
|  | 1745.62 | -1.208%
|  | 1736.09 | -1.747%
|  | 1760.8 | -0.3488%
|  | 1753.68 | -0.7517%
|  | 1755.4 | -0.6544%
|  | 1772.07 | 0.2891%
|  | 1791.95 | 1.414%
|  | 1767.18 | 0.01231%
|  | 1775.01 | 0.4554%
|  | 1756.04 | -0.6182%
| Total | 35339.25
| Average | 1766.9625
| Average | % Deviation from mean
| 1766.9625 | -5% | 1678.6144
|  | 5% | 1855.3106
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## Efa SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Tenofovir DF, Lamivudine & Efavirenz Tablets
Reference Substance: | Efavirenz
% age Purity: | 99.3
Weight (mg):
Standard Conc (mg/mL): | 0
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 106319119 | 87776.8 | 1.1 | 7.8
2 | 106634491 | 87582.9 | 1 | 7.8
3 | 106284869 | 87982.3 | 1.1 | 7.8
4 | 105969118 | 87552.6 | 1.2 | 7.8
5 | 106251700 | 87977.5 | 1 | 7.8
6 | 106624058 | 88236.2 | 1.1 | 7.8
Averages: | 106347225.8333 | 87851.3833 | 1.0833 | 7.8
RSD: | 0.2364%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Efavirenz
% age Purity: | 99.3
Weight (mg): | 0
Standard Conc (mg/mL): | 0
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 13900989 | 74086.6 | 1 | 6.9
2 | 13903091 | 74157 | 1 | 6.9
3 | 13914885 | 74306.4 | 1.1 | 6.9
4 | 13894782 | 74307.7 | 1.1 | 6.9
5 | 13924401 | 74561.7 | 1.1 | 6.9
6 | 13906607 | 74755.1 | 1 | 6.9
Averages: | 13907459.1667 | 74362.4167 | 1.05 | 6.9
RSD: | 0.0764%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2016-05-07
Reviewed By:

## TDF SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Tenofovir DF, Lamivudine & Efavirenz Tablets
Reference Substance: | Tenofovir DF
% age Purity: | 98.8
Weight (mg):
Standard Conc (mg/mL): | 0
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 25861612 | 111157.6 | 1.1 | 6.7
2 | 25952947 | 110861.6 | 1.1 | 6.7
3 | 25852804 | 110771.1 | 1.1 | 6.7
4 | 25751017 | 110687.6 | 1 | 6.7
5 | 25851801 | 111104.5 | 1.1 | 6.7
6 | 25972151 | 111581.4 | 1.1 | 6.7
Averages: | 25873722 | 111027.3 | 1.0833 | 6.7
RSD: | 0.3096%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Tenofovir DF
% age Purity: | 98.8
Weight (mg): | 0
Standard Conc (mg/mL): | 0
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 3087001 | 35592.9 | 0.9 | 5.9
2 | 3088896 | 41405.1 | 0.8 | 5.9
3 | 3086337 | 42517.9 | 0.9 | 5.9
4 | 3088984 | 39679.1 | 0.9 | 5.9
5 | 3085042 | 39147.4 | 0.8 | 5.9
6 | 3085664 | 36095.8 | 0.9 | 5.9
Averages: | 3086987.3333 | 39073.0333 | 0.86667 | 5.9
RSD: | 0.0534%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2016-05-07
Reviewed By:

## Lam SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | Tenofovir DF, Lamivudine & Efavirenz Tablets
Reference Substance: | Lamivudine
% age Purity: | 100.4
Weight (mg):
Standard Conc (mg/mL): | 0
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 42704578 | 11448 | 1.1 | 2.2
2 | 42852755 | 11504.8 | 1.1 | 2.2
3 | 42704507 | 11574.8 | 1.1 | 2.2
4 | 42544195 | 11517.7 | 1.1 | 2.2
5 | 42710573 | 11542.6 | 1.1 | 2.2
6 | 42898640 | 11557.3 | 1.1 | 2.2
Averages: | 42735874.6667 | 11524.2 | 1.1 | 2.2
RSD: | 0.2951%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance: | Lamivudine
% age Purity: | 100.4
Weight (mg): | 0
Standard Conc (mg/mL): | 0
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 5147090 | 10931.7 | 1.1 | 2.1
2 | 5153880 | 10613 | 1.1 | 2.1
3 | 5152618 | 10579.5 | 1.1 | 2.1
4 | 5161717 | 10410.8 | 1.1 | 2.1
5 | 5150480 | 10551.8 | 1.1 | 2.1
6 | 5153826 | 10549.1 | 1 | 2.1
Averages: | 5153268.5 | 10605.9833 | 1.0833 | 2.1
RSD: | 0.09444%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | Bugigi |  |  | 2016-05-07
Reviewed By:

## Lam
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Tenofovir DF, Lamivudine & Efavirenz Tablets
Laboratory Ref No: | NDQD2016061054 | 29
Active Ingredient: | Lamivudine
Label Claim: | Lamivudine 300mg
Date Analysis Started: | 2016-07-01
Date Analysis Completed: | 2016-07-05
Analysis Data
Reference Substance: | Lamivudine
Code: | L3 7
% age Purity: | 100.4
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 100.4
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 5 | 1 | 42467769 | 57429274.1236 | 40307370 | 57726394.0786
2nd Dilution Standard Transfer Volume (mL): | 4 | 2 | 42746526 | 57806237.9375 | 40207029 | 57582690.2322 |  | 0.1028%
2nd Dilution Standard Final Volume (mL): | 50 | 3 | 42810514 | 57892769.0758 | 40240507 | 57630635.9111
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 42674936.3333 | 57709427.0456 | 40251635.3333 | 57646573.4073
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 13.81 |  | 13.04
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 13.81 |  | 13.04
Standard Dilution Factor | 312.5 | Purity correction (mg): | 13.8652 |  | 13.0922
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.044369 |  | 0.041895
|  | Desired Concentration (mg/mL): | 0.06
|  | Desired Weight as free base (mg): | 18.75
|  | Desired Weight as salt (mg): | 18.75
|  | Average Normalised Response: | 57678000.2265
|  | RSD: | 0.2881%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Lamivudine 300mg
Each Tablet contains | 300 | Lamivudine
Average Tablet Content Weight (mg): | 1766.9625
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 1772.31 | 1 | 56486451 | 292.9159 | 97.64%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 56429137 | 292.6187 | 97.54%
2nd Dilution Sample Transfer Volume (mL): | 10 |  |  | 3 | 57234238 | 296.7937 | 98.93%
2nd Dilution Sample Final Volume (mL): | 50 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1755.8 | 1 | 57106485 | 298.9157 | 99.64%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 56663955 | 296.5994 | 98.87%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 57852016 | 302.8181 | 100.9%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1755.4 | 1 | 55660253 | 291.412 | 97.14%
Desired Sample Weight (mg): | 1766.9625 |  |  | 2 | 56627770 | 296.4775 | 98.83%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 56483381 | 295.7216 | 98.57%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 296.0303 | 98.68%
[5 empty] | RSD: | 1.184% | 1.184%
[5 empty] | n: | 9 | 9
Comment: | The content of | Lamivudine |  | in the sample as a percentage of the stated label claim is |  | 98.68%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Lamivudine
Code: | L3 7
% age Purity: | 100.4
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 100.4
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 5155652 | 7436801.6709 | 4876530 | 7449542.3215
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 5157306 | 7439187.4933 | 4858860 | 7422549.0675 |  | 0.03954%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 5155786 | 7436994.9601 | 4876984 | 7450235.8664
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 5156248 | 7437661.3748 | 4870791.3333 | 7440775.7518
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 13.81 |  | 13.04
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 13.81 |  | 13.04
Standard Dilution Factor | 66.6667 | Purity correction: | 13.8652 |  | 13.0922
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.20798 |  | 0.19638
|  | Desired Concentration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 20
|  | Desired Weight as salt (mg): | 20
|  | Average Normalised Peak Area: | 7439218.5633
|  | RSD: | 0.1366%
|  | n: | 6
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 6590412 | 265.7703 | 88.59%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 6660320 | 268.5895 | 89.53%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 6736098 | 271.6454 | 90.55%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 6752287 | 272.2982 | 90.77%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 6752985 | 272.3264 | 90.78%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 6775090 | 273.2178 | 91.07%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 270.6413 | 90.21%
Sample Dilution Factor | 1000 |  | RSD: | 1.06% | 1.06%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Lamivudine |  | dissolved as a percentage of the stated label claim is |  | 90.21%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## TDF
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Tenofovir DF, Lamivudine & Efavirenz Tablets
Laboratory Ref No: | NDQD2016061054 | 29
Active Ingredient: | Tenofovir DF
Label Claim: | Tenofovir DF 300mg
Date Analysis Started: | 2016-07-01
Date Analysis Completed: | 2016-07-05
Analysis Data
Reference Substance: | Tenofovir DF
Code: | T11 8
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 50 | 1 | 25716256 | 28665858.1391 | 28472523 | 29227528.5166
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 25898186 | 28868655.1392 | 28413301 | 29166736.1276 |  | 1.346%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 25953789 | 28930635.69 | 28389916 | 29142731.0279
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 25856077 | 28821716.3227 | 28425246.6667 | 29178998.5574
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 13.62 |  | 14.79
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 13.62 |  | 14.79
Standard Dilution Factor | 250 | Purity correction (mg): | 13.4566 |  | 14.6125
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.053826 |  | 0.05845
|  | Desired Concentration (mg/mL): | 0.06
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Response: | 29000357.4401
|  | RSD: | 0.7454%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Tenofovir DF 300mg
Each Tablet contains | 300 | Tenofovir DF
Average Tablet Content Weight (mg): | 1766.9625
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 1772.31 | 1 | 29023647 | 299.335 | 99.78%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 28967544 | 298.7564 | 99.59%
2nd Dilution Sample Transfer Volume (mL): | 10 |  |  | 3 | 29432439 | 303.5511 | 101.2%
2nd Dilution Sample Final Volume (mL): | 50 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1755.8 | 1 | 29072372 | 302.657 | 100.9%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 28870750 | 300.558 | 100.2%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 29478254 | 306.8824 | 102.3%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1755.4 | 1 | 28506583 | 296.8345 | 98.94%
Desired Sample Weight (mg): | 1766.9625 |  |  | 2 | 28998921 | 301.9611 | 100.7%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 28952782 | 301.4806 | 100.5%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 301.3351 | 100.4%
[5 empty] | RSD: | 0.9771% | 0.9771%
[5 empty] | n: | 9 | 9
Comment: | The content of | Tenofovir DF |  | in the sample as a percentage of the stated label claim is |  | 100.4%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Tenofovir DF
Code: | T11 8
% age Purity: | 98.8
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 98.8
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 3087125 | 3441211.9442 | 3382193 | 3471878.5671
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 3087142 | 3441230.8941 | 3383919 | 3473650.3355 |  | 0.9916%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 3085949 | 3439901.0594 | 3381408 | 3471072.7513
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 3086738.6667 | 3440781.2992 | 3382506.6667 | 3472200.5513
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 13.62 |  | 14.79
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 13.62 |  | 14.79
Standard Dilution Factor | 50 | Purity correction: | 13.4566 |  | 14.6125
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.26913 |  | 0.29225
|  | Desired Concentration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Peak Area: | 3456490.9253
|  | RSD: | 0.4987%
|  | n: | 6
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 3347498 | 290.5402 | 96.85%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 3390394 | 294.2632 | 98.09%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 3418666 | 296.7171 | 98.91%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 3434153 | 298.0612 | 99.35%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 3421848 | 296.9932 | 99%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 3441282 | 298.68 | 99.56%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 295.8758 | 98.63%
Sample Dilution Factor | 1000 |  | RSD: | 1.021% | 1.021%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Tenofovir DF |  | dissolved as a percentage of the stated label claim is |  | 98.63%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Efa
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | Tenofovir DF, Lamivudine & Efavirenz Tablets
Laboratory Ref No: | NDQD2016061054 | 29
Active Ingredient: | Efavirenz
Label Claim: | Efavirenz 600mg
Date Analysis Started: | 2016-07-01
Date Analysis Completed: | 2016-07-05
Analysis Data
Reference Substance: | Efavirenz
Code: | E15 3
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 4 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 105750699 | 79736576.9732 | 92157039 | 83041057.4451
2nd Dilution Standard Transfer Volume (mL): | 3 | 2 | 106285640 | 80139925.269 | 91875897 | 82787725.4237 |  | 2.937%
2nd Dilution Standard Final Volume (mL): | 50 | 3 | 106465038 | 80275192.2939 | 91835680 | 82751486.6052
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 106167125.6667 | 80050564.8454 | 91956205.3333 | 82860089.8247
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 33.39 |  | 27.94
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 33.39 |  | 27.94
Standard Dilution Factor | 208.3333 | Purity correction (mg): | 33.1563 |  | 27.7444
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.15915 |  | 0.13317
|  | Desired Concentration (mg/mL): | 0.12
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Response: | 81455327.335
|  | RSD: | 1.906%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Efavirenz 600mg
Each Tablet contains | 600 | Efavirenz
Average Tablet Content Weight (mg): | 1766.9625
Initial Sample dilution Volume (mL): | 100 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 5 | Assay Smp A | 1772.31 | 1 |  | - | -
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 84202752 | 618.3661 | 103.1%
2nd Dilution Sample Transfer Volume (mL): | 10 |  |  | 3 | 82987728 | 609.4433 | 101.6%
2nd Dilution Sample Final Volume (mL): | 50 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 1755.8 | 1 |  | - | -
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 85528031 | 634.0048 | 105.7%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 84113014 | 623.5155 | 103.9%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 5000 | Assay Smp C | 1755.4 | 1 | 85518902 | 634.0815 | 105.7%
Desired Sample Weight (mg): | 1766.9625 |  |  | 2 | 85672246 | 635.2185 | 105.9%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 83448911 | 618.7336 | 103.1%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 624.7662 | 104.1%
[5 empty] | RSD: | 1.594% | 1.594%
[5 empty] | n: | 7 | 7
Comment: | The content of | Efavirenz |  | in the sample as a percentage of the stated label claim is |  | 104.1%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Efavirenz
Code: | E15 3
% age Purity: | 99.3
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.3
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 5 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 10 | 1 | 13945415 | 12617898.5151 | 11885488 | 12851760.4621
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 13958514 | 12629750.5721 | 11840472 | 12803084.7284 |  | 1.649%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 13796122 | 12482817.2771 | 11855326 | 12819146.3365
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 13900017 | 12576822.1214 | 11860428.6667 | 12824663.8423
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 33.39 |  | 27.94
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 33.39 |  | 27.94
Standard Dilution Factor | 50 | Purity correction: | 33.1563 |  | 27.7444
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.66313 |  | 0.55489
|  | Desired Concentration (mg/mL): | 0.6
|  | Desired Weight as free base (mg): | 30
|  | Desired Weight as salt (mg): | 30
|  | Average Normalised Peak Area: | 12700742.9819
|  | RSD: | 1.15%
|  | n: | 6
Medium Volume (mL): | 1000 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 11900617 | 562.201 | 93.7%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 11829071 | 558.8211 | 93.14%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 11974006 | 565.668 | 94.28%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 11768392 | 555.9545 | 92.66%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 11736248 | 554.436 | 92.41%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 12521783 | 591.5457 | 98.59%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 564.771 | 94.13%
Sample Dilution Factor | 1000 |  | RSD: | 2.433% | 2.433%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Efavirenz |  | dissolved as a percentage of the stated label claim is |  | 94.13%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5