## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | PRIMOCEF 1G I G I.V
Reference Substance: | CEFOTAXIME
% age Purity: | 93.54
Weight (mg): | 10.16
Standard Conc (mg/mL): | 0.1016
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 33773743 | 3654.4 | 1.1 | 5.3
2 | 33724064 | 3646.6 | 1.1 | 5.3
3 | 33896068 | 3658 | 1 | 5.3
4 | 33614537 | 3669 | 1 | 5.3
5 | 33738588 | 3674.2 | 1.1 | 5.3
6 | 33731887 | 3658.3 | 1 | 5.3
Averages: | 33746481.1667 | 3660.0833 | 1.05 | 5.3
RSD: | 0.2693%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | MICHAEL
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | PRIMOCEF 1G I G I.V
Laboratory Ref No: |  | NDQD201607055
Active Ingredient: |  | CEFOTAXIME
Label Claim: |  | Each vial contains Sterile Cefotaxime sodium USP equivalent to Cefotaxime 1 g
Date Analysis Started: |  | 2016-07-28 12:58:56
Date Analysis Completed: |  | 2016-08-17 07:49:30
Analysis Data |  | Uniformity of weight
Capsule No. | Intact Capsule (mg) | Empty Shell (mg) | Capsule Content (mg) | % Deviation
1 | 15710.94 | 14639.71 | 1071.23 | -1.359%
2 | 14518.96 | 13437.87 | 1081.09 | -0.4508%
3 | 15706.37 | 14697.49 | 1008.88 | -7.1%
4 | 14731.03 | 13642.5 | 1088.53 | 0.2343%
5 | 14720.2 | 13634.49 | 1085.71 | -0.02541%
6 | 14769.71 | 13660.5 | 1109.21 | 2.139%
7 | 14747.3 | 13677.97 | 1069.33 | -1.534%
8 | 15474.76 | 14418.34 | 1056.42 | -2.723%
9 | 14648.62 | 13575.59 | 1073.03 | -1.193%
10 | 14579.53 | 13517.79 | 1061.74 | -2.233%
11 | 15883.65 | 14736.99 | 1146.66 | 5.587%
12 | 14722.86 | 13626.2 | 1096.66 | 0.9829%
13 | 14712.03 | 13611.86 | 1100.17 | 1.306%
14 | 14663.24 | 13538.47 | 1124.77 | 3.571%
15 | 14636.65 | 13550.78 | 1085.87 | -0.01068%
16 | 14644.39 | 13543.62 | 1100.77 | 1.361%
17 | 14732.43 | 13656.66 | 1075.77 | -0.9407%
18 | 14813.13 | 13716.35 | 1096.78 | 0.9939%
19 | 14733.11 | 13615.6 | 1117.51 | 2.903%
20 | 14672.84 | 13603.25 | 1069.59 | -1.51%
Total | 297821.75 | 276102.03 | 21719.72
Average | 14891.0875 | 13805.1015 | 1085.986
| Average | % Deviation from mean
| 1085.986 | -7.5% | 1004.537
|  | 7.5% | 1167.4349
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## CEFOTAXIME
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | PRIMOCEF 1G I G I.V
Laboratory Ref No: | NDQD201607055
Active Ingredient: | CEFOTAXIME
Label Claim: | Each vial contains Sterile Cefotaxime sodium USP equivalent to Cefotaxime 1 g
Date Analysis Started: | 2016-10-13
Date Analysis Completed: | 2016-10-14
Analysis Data
Reference Substance: | cefotaxime sodium
Code: | C21-2
% age Purity: | 93.54
% Water content: | 0 | If correction for water content is NOT needed, enter 0
% Purity corrected for water: | 93.54
Mwt of compound in free base form: | 455.47 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 477.45 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.95396 | free base
Each mg is equivalent to | 1667 | Insert the International Units equivalence
Initial Standard dilution | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 33825122 | 37309241.4098 | 40742005 | 38271255.4607
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 33613316 | 37075618.5663 | 40911608 | 38430573.0922
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 33730869 | 37205279.9835 | 40518803 | 38061589.2756
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 33723102.3333 | 37196713.3199 | 40724138.6667 | 38254472.6095
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 10.16 |  | 11.93
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 9.6923 |  | 11.3808
Standard Dilution Factor | 100 | Purity correction (mg): | 9.0662 |  | 10.6456
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 151.1327 |  | 177.462
|  | Desired Concetration (mg/mL): | 0.1
|  | Desired Weight as free base (iu): | 10
|  | Desired Weight as free base (mg): | 0.0059988
|  | Desired Weight as salt (mg) | 0.0062883
|  | Average Normalised Response: | 37725592.9647
|  | RSD: | 1.579%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each vial contains Sterile Cefotaxime sodium USP equivalent to Cefotaxime 1 g
Each Vial contains | 1000 | CEFOTAXIME
Average Vial Content Weight (mg): | 1085.986
Initial Sample dilution | 100 |  | Sample Weight (mg) | Injection | Response: | Determined Amt (iu) | % Assay
1st Dilution Sample transfer Volume (mL: | 5 | Assay Smp A | 115.34 | 1 |  | - | -
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 |  | - | -
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 |  | - | -
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 113.74 | 1 | 38043429 | 962.841 | 96.28%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 37872692 | 958.5198 | 95.85%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 37803892 | 956.7786 | 95.68%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 1000 | Assay Smp C | 105.96 | 1 | 35567085 | 966.2611 | 96.63%
Desired Sample Weight (mg): | 108.5986 |  |  | 2 | 35437880 | 962.7509 | 96.28%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 35380989 | 961.2053 | 96.12%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 96.14%
[6 empty] | RSD: | 0.3519%
[6 empty] | n: | 6
Comment: | The content of | CEFOTAXIME |  | in the sample as a percentage of the stated label claim is |  | 96.14%
| Name |  |  | Date |  | Signature
Analysed by: | MICHAEL
Reviewed By:
5