## Levamisole Hydrochloride
National Quality Control Laoboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | ZANIDE-L SUSPENSION 1000 mL
Laboratory Ref No: | NDQD2016061191 | 26
Active Ingredient: | Levamisole hydrochloride B.P
Date Analysis Started: | Each mL contains Oxyclozanide BP (Vet) 30 mg and Levamisole hydrochloride BP 15 mg
Date Analysis Completed: | 2016-07-19
Analysis Data | Standardisation of the Volumetric Solutions
Volumetric Solution: | Perchloric Acid
Reference Substance: | Potassium Hydrogen Phthalate
Molecular Weight: | 204.23
Target Concentration | 0.1
Reaction Ratio (Titrant:Standard) | 1 | : | 1
Standard | Standard Weight (mg) | mMoles of titrant | mL Titrant | Molarity (mM/mL) | Deviation from true Value | Correction Factor
A | 350.26 | 1.715 | 17.6 | 0.097445 | -2.555% | 0.97445
B | 350.18 | 1.7146 | 17.6 | 0.097422 | -2.578% | 0.97422
C | 350.09 | 1.7142 | 17.6 | 0.097397 | -2.603% | 0.97397
D |  | - |  | - | - | -
|  |  | Average : | 0.097422 | -2.578% | 0.97422
|  |  | RSD: | 0.02429%
|  |  | n: | 3
Analysis Data | Determination of Content of Active Ingredient in the Sample
Each | 1 | contains | 15 | Levamisole hydrochloride B.P
Each mL of | Perchloric Acid | is equivalent to | 24.08 | Levamisole hydrochloride B.P
Levamisole Content (mg)
Sample | Suspension Volume (mL) | Titre Vol. (mL) | Blank | Blank Correction | Corrected Titre | In sample | Per Label Claim | Percentage content
A | 25 | 15.5 | 0 | 15.5 | 15.1003 | 363.6162 | 14.5446 | 96.96%
B | 25 | 15.4 | 0 | 15.4 | 15.0029 | 361.2703 | 14.4508 | 96.34%
C | 25 | 15.4 | 0 | 15.4 | 15.0029 | 361.2703 | 14.4508 | 96.34%
D |  |  |  | - | - | - | - | -
|  | Average : | 0 |  | Average : | 362.0522 | 14.4821 | 96.55%
|  | RSD: | - |  | RSD: |  | 0.3741% | 0.3741%
|  | n: | 3 |  | n: | 3 | 3 | 3
| Name |  |  | Date |  | Signature
Analysed by: | RUTTO/JOYFRIDA
Reviewed By:
0

## Levamisole Hydrochloride 1
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Relative Density Test Report
#REF!
Sample Name: | ZANIDE - L SUSPENSION 1000 mL
Laboratory Ref No: | NDQD2016061191
Active Ingredient: | Levamisole Hydrochloride
Label Claim: | Each mL contains Oxyclozanide BP (Vet) 30 mg and Levamisole hydrochloride BP 15 mg
Date Analysis Started: | 19/07/2016
Date Analysis Completed: | 19/07/2016
| Pyknometer Mass (g) | Pyknometer + Water (g) | Pyknometer + Sample (g)
| 9.9668 | 15.7792 | 16.0338
|  | 15.7631 | 16.0257
|  | 15.7648 | 16.0283
| 9.9668 | 15.769 | 16.0293
| Mass of Water (g): | 5.8023
| Mass of Sample (g): | 6.0625
| Relative Density of Sample: | 1.0448
| Name |  | Date | Signature
Analysed by:
Reviewed By:
0