## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | MYDAWA CO-AMOXACILLIN 1000 TABLETS
Laboratory Ref No: |  | NDQD2016061218
Active Ingredient: |  | Amoxicillin & Clavulanate Potassium
Label Claim: |  | Each tablet contains Amoxicillin Trihydrate 875 mg Clavulanate Potassium 125 mg
Date Analysis Started: |  | 2016-06-24 13:14:58
Date Analysis Completed: |  | #REF!
Analysis Data |  | Uniformity of weight
|  | Tablet weight (mg) | % Deviation
|  | 1466.95 | -0.182%
|  | 1451.6 | -1.226%
|  | 1478.63 | 0.6128%
|  | 1486.2 | 1.128%
|  | 1466.63 | -0.2037%
|  | 1454.03 | -1.061%
|  | 1476.67 | 0.4794%
|  | 1473.69 | 0.2767%
|  | 1455.85 | -0.9372%
|  | 1452.28 | -1.18%
|  | 1478.89 | 0.6305%
|  | 1481.44 | 0.804%
|  | 1472.72 | 0.2107%
|  | 1481.35 | 0.7979%
|  | 1464.47 | -0.3507%
|  | 1471.15 | 0.1038%
|  | 1465.23 | -0.299%
|  | 1479.26 | 0.6557%
|  | 1476.09 | 0.44%
|  | 1459.35 | -0.6991%
| Total | 29392.48
| Average | 1469.624
| Average | % Deviation from mean
| 1469.624 | -5% | 1396.1428
|  | 5% | 1543.1052
| Name |  | Date | Signature
Analysed by:
Reviewed By:

## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | MYDAWA CO-AMOXACILLIN 1000 TABLETS
Reference Substance: | Amoxicillin
% age Purity: | 87.84
Weight (mg): | 21.99
Standard Conc (mg/mL): | 0.5
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min) | Peak Resolution (USP)
1 | 85087902 | 5287 | 1.1 | 6 | 8.9
2 | 85865599 | 5265.2 | 1.1 | 6 | 8.9
3 | 85069413 | 5245.1 | 1.1 | 6 | 8.9
4 | 85393505 | 5196.8 | 1.1 | 6 | 8.9
5 | 85488979 | 5202.4 | 1.1 | 6 | 8.9
6 | 85501760 | 5182.2 | 1.1 | 6 | 8.8
Averages: | 85401193 | 5229.7833 | 1.1 | 6 | 8.8833
RSD: | 0.3481%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
The Peak Resolution (USP) between Amoxicillin & Clavulanic Lithium is greater than 3.5
Analysis Data | CLAVULANIC
Reference Substance: | Clavulanic lithium
% age Purity: | 96.96
Weight (mg): | 9.46
Standard Conc (mg/mL): | 0.125
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 30600135 | 6720.3 | 1.3 | 3.7
2 | 30793764 | 6737.5 | 1.3 | 3.7
3 | 30568266 | 6705.9 | 1.3 | 3.7
4 | 30679684 | 6663.8 | 1.3 | 3.7
5 | 30704248 | 6630.5 | 1.3 | 3.7
6 | 30689325 | 6616 | 1.3 | 3.7
Averages: | 30672570.3333 | 6679 | 1.3 | 3.7
RSD: | 0.2614%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:

## amoxicillin Trihydrate
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | MYDAWA CO-AMOXACILLIN 1000 TABLETS
Laboratory Ref No: | NDQD2016061218 | 29
Active Ingredient: | Amoxicillin & Clavulanate Potassium
Label Claim: | Each tablet contains Amoxicillin Trihydrate 875 mg Clavulanate Potassium 125 mg
Date Analysis Started: | 2016-06-24 13:14:58
Date Analysis Completed:
Analysis Data
Reference Substance: | Amoxicillin Trihydrate
Code: | wrs A1-3
% age Purity: | 87.84
% Water content: |  | If correction for water content is not needed please enter 0
% Purity corrected for water: | 87.84
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 86184319 | 111545153.7729 | 82125641 | 109324815.4758
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 84791685 | 109742719.4614 | 80836453 | 107608661.5621 |  | 1.885%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 85167634 | 110229296.2483 | 81247108 | 108155321.3087
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 85381212.6667 | 110505723.1609 | 81403067.3333 | 108362932.7822
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 21.99 |  | 21.38
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 21.99 |  | 21.38
Standard Dilution Factor | 50 | Purity correction (mg): | 19.316 |  | 18.7802
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.38632 |  | 0.3756
|  | Desired Concentration (mg/mL): | 0.5
|  | Desired Weight as free base (mg): | 25
|  | Desired Weight as salt (mg): | 25
|  | Average Normalised Response: | 109434327.9715
|  | RSD: | 1.303%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains Amoxicillin Trihydrate 875 mg Clavulanate Potassium 125 mg
Each Tablet contains | 875 | Amoxicillin & Clavulanate Potassium
Average Tablet Content Weight (mg): | 1469.624
Initial Sample dilution Volume (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 135.59 | 1 | 367205051 | 909.2298 | 103.9%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 369717003 | 915.4496 | 104.6%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 368418639 | 912.2347 | 104.3%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 127.22 | 1 | 353615594 | 933.187 | 106.6%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 355232552 | 937.4541 | 107.1%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 356389125 | 940.5063 | 107.5%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 50 | Assay Smp C | 154.77 | 1 | 421395325 | 914.104 | 104.5%
Desired Sample Weight (mg): | 41.9893 |  |  | 2 | 422410385 | 916.3059 | 104.7%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 421675451 | 914.7116 | 104.5%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 921.4648 | 105.3%
[5 empty] | RSD: | 1.303% | 1.303%
[5 empty] | n: | 9 | 9
Comment: | The content of | Amoxicillin & Clavulanate Potassium |  | in the sample as a percentage of the stated label claim is |  | 105.3%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Amoxicillin Trihydrate
Code: | wrs A1-3
% age Purity: | 87.84
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 87.84
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 20 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 10 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 86184319 | 216893354.5584 | 82125641 | 212576030.0918
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 84791685 | 213388621.175 | 80836453 | 209239064.1486 |  | 1.885%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 85167634 | 214334742.7049 | 81247108 | 210302013.6559
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 85381212.6667 | 214872239.4794 | 81403067.3333 | 210705702.6321
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 21.99 |  | 21.38
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 21.99 |  | 21.38
Standard Dilution Factor | 50 | Purity correction: | 19.316 |  | 18.7802
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.38632 |  | 0.3756
|  | Desired Concentration (mg/mL): | 0.97222
|  | Desired Weight as free base (mg): | 48.6111
|  | Desired Weight as salt (mg): | 48.6111
|  | Average Normalised Peak Area: | 212788971.0558
|  | RSD: | 1.303%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 212379533 | 873.3164 | 99.81%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 193600281 | 796.095 | 90.98%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 212341839 | 873.1614 | 99.79%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 193100861 | 794.0414 | 90.75%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 214864953 | 883.5366 | 101%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 193095747 | 794.0204 | 90.75%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 835.6952 | 95.51%
Sample Dilution Factor | 900 |  | RSD: | 5.391% | 5.391%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Amoxicillin & Clavulanate Potassium |  | dissolved as a percentage of the stated label claim is |  | 95.51%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5

## Clavulanate lithium
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | MYDAWA CO-AMOXACILLIN 1000 TABLETS
Laboratory Ref No: | NDQD2016061218 | 29
Active Ingredient: | Amoxicillin & Clavulanate Potassium
Label Claim: | Each tablet contains Amoxicillin Trihydrate 875 mg Clavulanate Potassium 125 mg
Date Analysis Started: | 2016-06-24 13:14:58
Date Analysis Completed:
Analysis Data
Reference Substance: | Clavulanic Lithium
Code: | wrs C64-2
% age Purity: | 96.96
% Water content: |  | If correction for water content is not needed please enter 0
% Purity corrected for water: | 96.96
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 10 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 2 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 20067229 | 21877800.7888 | 26912754 | 22187492.5802
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 20045100 | 21853675.1931 | 27020879 | 22276633.3138 |  | 2.104%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 20126594 | 21942522.0138 | 26994552 | 22254928.7303
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 20079641 | 21891332.6652 | 26976061.6667 | 22239684.8748
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 9.46 |  | 12.51
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 9.46 |  | 12.51
Standard Dilution Factor | 125 | Purity correction (mg): | 9.1724 |  | 12.1297
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.073379 |  | 0.097038
|  | Desired Concentration (mg/mL): | 0.08
|  | Desired Weight as free base (mg): | 10
|  | Desired Weight as salt (mg): | 10
|  | Average Normalised Response: | 22065508.77
|  | RSD: | 0.8848%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each tablet contains Amoxicillin Trihydrate 875 mg Clavulanate Potassium 125 mg
Each Tablet contains | 125 | Amoxicillin & Clavulanate Potassium
Average Tablet Content Weight (mg): | 1469.624
Initial Sample dilution Volume (mL): | 50 |  | Powder Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL): | 1 | Assay Smp A | 135.59 | 1 | 66934207 | 131.5142 | 105.2%
1st Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 67288491 | 132.2103 | 105.8%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 66854370 | 131.3574 | 105.1%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 127.22 | 1 | 62765742 | 131.4376 | 105.2%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 62834886 | 131.5824 | 105.3%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 63018473 | 131.9668 | 105.6%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 50 | Assay Smp C | 154.77 | 1 | 76215675 | 131.1927 | 105%
Desired Sample Weight (mg): | 47.028 |  |  | 2 | 76313758 | 131.3616 | 105.1%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 76065078 | 130.9335 | 104.7%
|  |  |  | 4 |  | - | -
[5 empty] | Average: | 131.5063 | 105.2%
[5 empty] | RSD: | 0.293% | 0.293%
[5 empty] | n: | 9 | 9
Comment: | The content of | Amoxicillin & Clavulanate Potassium |  | in the sample as a percentage of the stated label claim is |  | 105.2%
DISSOLUTION: | Determination of Active Ingredient Dissolved
Reference Substance: | Clavulanic Lithium
Code: | wrs C64-2
% age Purity: | 96.96
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 96.96
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt convertion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt convertion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution volume (mL): | 10 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Injection | Response: | Normalised Response: | Response: | Normalised Response: |  | Std Response Deviation
1st Dilution Standard Final Volume (mL): | 25 | 1 | 30878649 | 38963755.1146 | 40793568 | 38924947.3175
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 30342598 | 38287347.2869 | 39865613 | 38039498.9182 |  | 0.4105%
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 30385200 | 38341103.9747 | 40106115 | 38268984.3038
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 30535482.3333 | 38530735.4587 | 40255098.6667 | 38411143.5132
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 9.46 |  | 12.51
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 9.46 |  | 12.51
Standard Dilution Factor | 83.3333 | Purity correction: | 9.1724 |  | 12.1297
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Conc (mg/mL): | 0.11007 |  | 0.14556
|  | Desired Concentration (mg/mL): | 0.13889
|  | Desired Weight as free base (mg): | 11.5741
|  | Desired Weight as salt (mg): | 11.5741
|  | Average Normalised Peak Area: | 38470939.486
|  | RSD: | 0.9908%
|  | n: | 6
Medium Volume (mL): | 900 | Capsule No. | Response: | Amt Released (mg): | %age Released:
1st Dilution Sample transfer Volume (mL: | 1 | 1 | 36008404 | 116.9987 | 93.6%
1st Dilution Sample Final Volume (mL): | 1 | 2 | 37148147 | 120.702 | 96.56%
2nd Dilution Sample Transfer Volume (mL): | 1 | 3 | 35973615 | 116.8857 | 93.51%
2nd Dilution Sample Final Volume (mL): | 1 | 4 | 36933505 | 120.0046 | 96%
3rd Dilution Sample Transfer Volume (mL: | 1 | 5 | 36310906 | 117.9816 | 94.39%
3rd Dilution Sample Final Volume (mL): | 1 | 6 | 36950913 | 120.0611 | 96.05%
4th Dilution Sample Transfer Volume (mL): | 1
4th Dilution Sample Final Volume (mL): | 1 |  | Average: | 118.7723 | 95.02%
Sample Dilution Factor | 900 |  | RSD: | 1.421% | 1.421%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  | n: | 6 | 6
Comment: | The amount of | Amoxicillin & Clavulanate Potassium |  | dissolved as a percentage of the stated label claim is |  | 95.02%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5