## RD
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Relative Density Test Report
COLDIFED -P SYRUP 120 ML
NDQD201607011
Sample Name: | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Laboratory Ref No: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Active Ingredient: | 2016-07-05 12:39:40
Label Claim: | 0
Date Analysis Started: | 0
Date Analysis Completed: | 0
|  | Pyknometer + Water (g) | Pyknometer + Sample (g)
| 9.478 | 15.3811 | 16.5094
|  | 15.3808 | 16.4944
|  | 15.3572 | 16.4825
| 9.478 | 15.373 | 16.4954
| Mass of Water (g): | 5.895
| Mass of Sample (g): | 7.0174
| Relative Density of Sample: | 1.1904
| Name |  | Date | Signature
Analysed by:
Reviewed By:
0

## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | COLDIFED -P SYRUP 120 ML
Reference Substance: | Paracetamol (Acetominophen)
% age Purity: | 99.4
Weight (mg): | 31.26
Standard Conc (mg/mL): | 0.01563
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 1354845 | 6080.5 | 1.1 | 3.9
2 | 1355404 | 6117.1 | 1.2 | 3.9
3 | 1354025 | 6131.5 | 1.1 | 3.9
4 | 1358152 | 6056.2 | 1.1 | 3.9
5 | 1354124 | 6114.3 | 1.2 | 3.9
6 | 1356088 | 6121.5 | 1.1 | 3.9
Averages: | 1355439.6667 | 6103.5167 | 1.1333 | 3.9
RSD: | 0.1136%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Assay
Reference Substance: | Pseudoephedrine HCl
% age Purity: | 100.5
Weight (mg): | 15.65
Standard Conc (mg/mL): | 0.313
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 537320 | 5503.4 | 1.7 | 7.1
2 | 539618 | 5333.8 | 1.7 | 7.1
3 | 539761 | 5448.5 | 1.7 | 7.1
4 | 542587 | 5461.5 | 1.7 | 7.2
5 | 540937 | 5469.3 | 1.7 | 7.2
6 | 535556 | 5400 | 1.6 | 7.1
Averages: | 539296.5 | 5436.0833 | 1.6833 | 7.1333
RSD: | 0.4672%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Assay
Reference Substance: | Triprolidine HCl
% age Purity: | 94.6
Weight (mg): | 12.39
Standard Conc (mg/mL): | 0.014868
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min) | Resolution
1 | 634660 | 7663.9 | 1.2 | 20 | 20
2 | 630210 | 7352.2 | 1.1 | 20 | 19.7
3 | 662440 | 7350.8 | 1.2 | 20.3 | 19.7
4 | 636392 | 7549.9 | 1.1 | 19.9 | 19.9
5 | 644636 | 7547.2 | 1.1 | 20.4 | 19.9
6 | 636660 | 7451.3 | 1.1 | 20 | 19.7
Averages: | 640833 | 7485.8833 | 1.1333 | 20.1 | 19.8167
RSD: | 1.806%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date |  | Signature
Analysed by: |  | Dr. Sarah Mwangi |  | 2nd August 2016
Reviewed By:

## Paracetamol
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | COLDIFED -P SYRUP 120 ML
Laboratory Ref No: | NDQD201607011 | 25
Active Ingredient: | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Label Claim: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Date Analysis Started: | 2016-07-25
Date Analysis Completed: | 2016-07-27
Analysis Data
Reference Substance: | Paracetamol Acetominophen
Code: | P49-2
% age Purity: | 99.4
% Water content: |  | If correction for water content is not needed please enter 0
% Purity corrected for water: | 99.4
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 5 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 100 | 1 | 1357729 | 1048694.4701 | 1285460 | 1042218.3876
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 1354681 | 1046340.2295 | 1283521 | 1040646.2955
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 1354817 | 1046445.2743 | 1283473 | 1040607.3783
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 1355742.3333 | 1047159.9913 | 1284151.3333 | 1041157.3538
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 31.26 |  | 29.78
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 31.26 |  | 29.78
Standard Dilution Factor | 2000 | Purity correction (mg): | 31.0724 |  | 29.6013
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.015536 |  | 0.014801
|  | Desired Concentration (mg/mL): | 0.012
|  | Desired Weight as free base (mg): | 24
|  | Desired Weight as salt (mg): | 24
|  | Average Normalised Response: | 1044158.6725
|  | RSD: | 0.3297%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Each | 5 | contains | 120 | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Relative Density of sample: | 1.1904
Each | 5 | is equivalent to | 5.952
Initial Sample dilution (mL): | 100 |  | Powder Weight (g) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 1 | Assay Smp A | 5.8943 | 1 | 1045049 | 121.2778 | 101.1%
1st Dilution Sample Final Volume (mL): | 100 |  |  | 2 | 1048236 | 121.6477 | 101.4%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 1043167 | 121.0594 | 100.9%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 6.3712 | 1 | 1148500 | 123.3079 | 102.8%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 1148736 | 123.3333 | 102.8%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 1147763 | 123.2288 | 102.7%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 10000 | Assay Smp C | 5.8866 | 1 | 1044521 | 121.3749 | 101.1%
Desired Powder Weight (g): | 5.952 |  |  | 2 | 1042457 | 121.1351 | 100.9%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 1042384 | 121.1266 | 100.9%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 101.6%
[6 empty] | RSD: | 0.8404%
[6 empty] | n: | 9
Comment: | The content of | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl |  | in the sample as a percentage of the stated label claim is |  | 101.6%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## Pseudoephedrine HCL 1
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | COLDIFED -P SYRUP 120 ML
Laboratory Ref No: | NDQD201607011 | 25
Active Ingredient: | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Label Claim: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Date Analysis Started: | 2016-07-25
Date Analysis Completed: | 2016-08-02
Analysis Data
Reference Substance: | Pseudoephedrine HCl
Code: | P58-1
% age Purity: | 100.5
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 100.5
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 535556 | 510758.6667 | 661179 | 507109.1102
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 530516 | 505952.029 | 649125 | 497863.969
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 524146 | 499876.973 | 647524 | 496636.0387
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 530072.6667 | 505529.2229 | 652609.3333 | 500536.3726
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 15.65 |  | 19.46
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 15.65 |  | 19.46
Standard Dilution Factor | 50 | Purity correction (mg): | 15.7283 |  | 19.5573
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.31457 |  | 0.39115
|  | Desired Concentration (mg/mL): | 0.3
|  | Desired Weight as free base (mg): | 15
|  | Desired Weight as salt (mg): | 15
|  | Average Normalised Response: | 503032.7978
|  | RSD: | 1.133%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Each | 5 | contains | 30 | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Relative Density of sample: | 1.1904
Each | 5 | is equivalent to | 5.952
Initial Sample dilution (mL): | 50 |  | Powder Weight (g) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 15 | Assay Smp A | 6.4032 | 1 | 654010 | 30.213 | 100.7%
1st Dilution Sample Final Volume (mL): | 25 |  |  | 2 | 647476 | 29.9111 | 99.7%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 651720 | 30.1072 | 100.4%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 6.4129 | 1 | 654809 | 30.2045 | 100.7%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 647731 | 29.878 | 99.59%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 654306 | 30.1813 | 100.6%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 83.3333 | Assay Smp C | 6.4134 | 1 | 649278 | 29.9468 | 99.82%
Desired Powder Weight (g): | 4.96 |  |  | 2 | 644616 | 29.7317 | 99.11%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 652466 | 30.0938 | 100.3%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 100.1%
[6 empty] | RSD: | 0.5646%
[6 empty] | n: | 9
Comment: | The content of | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl |  | in the sample as a percentage of the stated label claim is |  | 100.1%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0

## Tripolidine HCL
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | COLDIFED -P SYRUP 120 ML
Laboratory Ref No: | NDQD201607011 | 25
Active Ingredient: | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Label Claim: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Date Analysis Started: | 2016-07-25
Date Analysis Completed: | 2016-08-02
Analysis Data
Reference Substance: | Triprolidine HCl
Code: | T35-1
% age Purity: | 94.6
% Water content: | 0 | If correction for water content is not needed please enter 0
% Purity corrected for water: | 94.6
Mwt of compound in free base form: | 1 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 1 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 1 | free base
Initial Standard dilution (mL): | 50 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 3 | Inj | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 50 | 1 | 636660 | 543181.6902 | 715109 | 562447.3032
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 651603 | 555930.6677 | 702659 | 552655.1331
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 641498 | 547309.3455 | 719862 | 566185.6312
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 643253.6667 | 548807.2345 | 712543.3333 | 560429.3559
4th Dilution Standard Transfer Volume (mL): | 1 | Mass of RS (mg): | 12.39 |  | 13.44
4th Dilution Standard Final Volume (mL): | 1 | Mass of WRS as free base (mg): | 12.39 |  | 13.44
Standard Dilution Factor | 833.3333 | Purity correction (mg): | 11.7209 |  | 12.7142
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.014065 |  | 0.015257
|  | Desired Concentration (mg/mL): | 0.012
|  | Desired Weight as free base (mg): | 10
|  | Desired Weight as salt (mg): | 10
|  | Average Normalised Response: | 554618.2952
|  | RSD: | 1.582%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each 5 mL contains Paracetamol 120 mg, Pseudoephedrine Hcl 30 mg, Triprolidine Hcl 1.25 mg
Each | 5 | contains | 1.25 | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl
Relative Density of sample: | 1.1904
Each | 5 | is equivalent to | 5.952
Initial Sample dilution (mL): | 50 |  | Powder Weight (g) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 15 | Assay Smp A | 6.4032 | 1 | 701916 | 1.1764 | 94.11%
1st Dilution Sample Final Volume (mL): | 25 |  |  | 2 | 686042 | 1.1498 | 91.98%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 695525 | 1.1657 | 93.26%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 6.4129 | 1 | 698537 | 1.169 | 93.52%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 683322 | 1.1435 | 91.48%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 691144 | 1.1566 | 92.53%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 83.3333 | Assay Smp C | 6.4134 | 1 | 685022 | 1.1463 | 91.7%
Desired Powder Weight (g): | 4.7616 |  |  | 2 | 678438 | 1.1352 | 90.82%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 698700 | 1.1692 | 93.53%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 92.55%
[6 empty] | RSD: | 1.209%
[6 empty] | n: | 9
Comment: | The content of | Paracetamol, Pseudoephedrine Hcl, Triprolidine Hcl |  | in the sample as a percentage of the stated label claim is |  | 92.55%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
0