## SST
HPLC System Suitability Report
Analysis Data | Assay
Sample(s) | PRIMOCEF 1G I G I.V
Reference Substance: | CEFOTAXIME
% age Purity: | 93.54
Weight (mg): | 10.16
Standard Conc (mg/mL): | 0.1016
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1 | 33773743 | 3654.4 | 1.1 | 5.3
2 | 33724064 | 3646.6 | 1.1 | 5.3
3 | 33896068 | 3658 | 1 | 5.3
4 | 33614537 | 3669 | 1 | 5.3
5 | 33738588 | 3674.2 | 1.1 | 5.3
6 | 33731887 | 3658.3 | 1 | 5.3
Averages: | 33746481.1667 | 3660.0833 | 1.05 | 5.3
RSD: | 0.2693%
n: | 6
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
Analysis Data | Dissolution
Reference Substance:
% age Purity:
Weight (mg):
Standard Conc (mg/mL):
Injection Number | Peak Areas | Theoretical Plates (USP) | Tailing Factor (Assym) | Retention Times (min)
1
2
3
4
5
6
Averages: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
RSD: | #DIV/0!
n: | 0
Summary: | The RSD of the peak areas for six replicate injections of SST Std is less than 2.0%.
The number of Theoretical Plates (USP) for all peaks is greater than 2000
The Assymetry of all peaks were below 2.0
| Name |  |  | Date | Signature
Analysed by: | MICHAEL
Reviewed By:

## Uniformity
Please enter the required information in the cells highlighted in green
Uniformity of Weight Test Report
Sample Name: |  | PRIMOCEF 0.5 G I.M/I.V
Laboratory Ref No: |  | NDQD201607054
Active Ingredient: |  | CEFOTAXIME
Label Claim: |  | Each vial contains Sterile Cefotaxime sodium USP equivalent to Cefotaxime 0.5 g
Date Analysis Started: |  | 2016-07-28 12:42:01
Date Analysis Completed: |  | 2016-08-16 13:07:29
Analysis Data |  | Uniformity of weight
Capsule No. | Intact Capsule (mg) | Empty Shell (mg) | Capsule Content (mg) | % Deviation
1 | 15085.88 | 14502.56 | 583.32 | 0.119%
2 | 15108.68 | 14535.09 | 573.59 | -1.551%
3 | 15232.33 | 14658.56 | 573.77 | -1.52%
4 | 15154.37 | 14556.23 | 598.14 | 2.663%
5 | 14973.21 | 14398.95 | 574.26 | -1.436%
6 | 14920.53 | 14333.45 | 587.08 | 0.7644%
7 | 15328.46 | 14737.16 | 591.3 | 1.489%
8 | 15207.53 | 14643.82 | 563.71 | -3.247%
9 | 15116.66 | 14513.19 | 603.47 | 3.578%
10 | 15275.92 | 14683.42 | 592.5 | 1.695%
11 | 15038.72 | 14467.01 | 571.71 | -1.874%
12 | 15133.3 | 14541.03 | 592.27 | 1.655%
13 | 15051.39 | 14475.75 | 575.64 | -1.199%
14 | 15231.14 | 14634.82 | 596.32 | 2.35%
15 | 15004.65 | 14419.32 | 585.33 | 0.464%
16 | 15251.61 | 14667.42 | 584.19 | 0.2684%
17 | 15166.43 | 14565.23 | 601.2 | 3.188%
18 | 15178.18 | 14626.65 | 551.53 | -5.337%
19 | 15219.47 | 14651.69 | 567.78 | -2.548%
20 | 15151.4 | 14565.98 | 585.42 | 0.4795%
Total | 302829.86 | 291177.33 | 11652.53
Average | 15141.493 | 14558.8665 | 582.6265
| Average | % Deviation from mean
| 582.6265 | -7.5% | 538.9295
|  | 7.5% | 626.3235
| Name |  |  | Date | Signature
Analysed by:
Reviewed By:

## CEFOTAXIME
National Quality Control Laboratory
Laboratory Data Calculation Spreadsheet
Please enter the required information in the cells highlighted in green
Analysis Report
Sample Name: | PRIMOCEF 0.5 G I.M/I.V
Laboratory Ref No: | NDQD201607054 | 8
Active Ingredient: | CEFOTAXIME
Label Claim: | Each vial contains Sterile Cefotaxime sodium USP equivalent to Cefotaxime 0.5 g
Date Analysis Started: | 2016-10-13
Date Analysis Completed: | 2016-10-14
Analysis Data
Reference Substance: | CEFOTAXIME SODIUM
Code: | C21-2
% age Purity: | 93.54
% Water content: | 0 | If correction for water content is NOT needed, enter 0
% Purity corrected for water: | 93.54
Mwt of compound in free base form: | 455.47 | Enter molecular mass of compound in free base form. If salt conversion is NOT needed, enter 1.
Mwt of compound in salt form: | 477.45 | Enter molecular mass of compound in salt form. If salt conversion is NOT needed, enter 1.
1 mg of salt is equivalent to | 0.95396 | free base
Initial Standard dilution (mL): | 100 |  | Standard A |  | Standard B
1st Dilution Standard transfer Volume (mL: | 1 | Injection | Response: | Normalised Response: | Response: | Normalised Response:
1st Dilution Standard Final Volume (mL): | 1 | 1 | 33825122 | 37309241.4098 | 40742005 | 38271255.4607
2nd Dilution Standard Transfer Volume (mL): | 1 | 2 | 33613316 | 37075618.5663 | 40911608 | 38430573.0922
2nd Dilution Standard Final Volume (mL): | 1 | 3 | 33730869 | 37205279.9835 | 40518803 | 38061589.2756
3rd Dilution Standard Volume (mL: | 1 | 4 |  | - |  | -
3rd Dilution Standard Final Volume (mL): | 1 | Average: | 33723102.3333 | 37196713.3199 | 40724138.6667 | 38254472.6095
4th Dilution Standard Transfer Volume (mL): | 1 | Amt of RS (mg): | 10.16 |  | 11.93
4th Dilution Standard Final Volume (mL): | 1 | Amt of RS as free base (mg): | 9.6923 |  | 11.3808
Standard Dilution Factor | 100 | Purity correction (mg): | 9.0662 |  | 10.6456
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  | Concentration (mg/mL): | 0.090662 |  | 0.10646
|  | Desired Concetration (mg/mL): | 0.1
|  | Desired Weight as free base (mg): | 10
|  | Desired Weight as salt (mg): | 10.4826
|  | Average Normalised Response: | 37725592.9647
|  | RSD: | 1.579%
|  | n: | 6
Analysis Data | Determination of Content of Active Ingredient in the Sample
Label Claim: | Each vial contains Sterile Cefotaxime sodium USP equivalent to Cefotaxime 0.5 g
Each Vial contains | 500 | CEFOTAXIME
Average Vial Content Weight (mg): | 582.6265
Initial Sample dilution (mL): | 50 |  | Sample Weight (mg) | Injection | Response: | Determined Amt (mg) | % Assay
1st Dilution Sample transfer Volume (mL: | 5 | Assay Smp A | 63.19 | 1 | 42418254 | 518.3566 | 103.7%
1st Dilution Sample Final Volume (mL): | 50 |  |  | 2 | 42216061 | 515.8857 | 103.2%
2nd Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 41887954 | 511.8762 | 102.4%
2nd Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
3rd Dilution Sample Transfer Volume (mL: | 1 | Assay Smp B | 59.73 | 1 | 40025142 | 517.4454 | 103.5%
3rd Dilution Sample Final Volume (mL): | 1 |  |  | 2 | 39948918 | 516.46 | 103.3%
4th Dilution Sample Transfer Volume (mL): | 1 |  |  | 3 | 39593921 | 511.8706 | 102.4%
4th Dilution Sample Final Volume (mL): | 1 |  |  | 4 |  | - | -
Sample Dilution Factor | 500 | Assay Smp C | 59.4 | 1 | 38748892 | 503.7291 | 100.7%
Desired Sample Weight (mg): | 58.2626 |  |  | 2 | 38653262 | 502.4859 | 100.5%
If there are no serial dilutions, or only one dilution, enter 1 in all boxes not used. |  |  |  | 3 | 38299286 | 497.8843 | 99.58%
|  |  |  | 4 |  | - | -
[6 empty] | Average: | 102.1%
[6 empty] | RSD: | 1.464%
[6 empty] | n: | 9
Comment: | The content of | CEFOTAXIME |  | in the sample as a percentage of the stated label claim is |  | 102.1%
| Name |  |  | Date |  | Signature
Analysed by:
Reviewed By:
5